## Instructions
Template for Entering Candidate Method Test
Data and Calculating Results Related to a U.S.
EPA Equivalent Method Designation Application
for PM2.5 and PM10-2.5 [40 CFR 53.35]
Instructions
This suite of spreadsheets is intended for entering and analyzing field test data obtained from the Comparability tests required for Class II or III candidate equivalent methods for PM2.5 or PM10-2.5, according to new regulatory provisions of 40 CFR Part 53 that became effective on December 18, 2006. (See Section 53.35 of Part 53, Subpart C (U.S. Federal Register, Volume 71, October 17, 2006, page 61278).
This spreadsheet is designed to run on Excel 2003. It appears to run properly on Excel 2000, but the Precision and Regression sheets must be unprotected (Tools>Protection>Unprotect Sheet) to make the data filters work. Also, for better viewing, the zoom should be set to 75%.
1. Enter applicant, candidate method, and site identification information in the "Title" tab. It is very important to properly identify the PM type (i.e., PM2.5 or PM10-2.5) and the site category (A, B, C, or D) because this information is used to determine some of the pass-fail criteria for certain tests.
2. The test data should be in the form of daily "data sets" consisting of 22-25 hour integrated ambient PM concentration measurements, each set containing nominally 3 FRM (FRM subset) and 3 candidate FEM (FEM subset) measurements, in μg/m3.
3. Enter the FRM data first, at the "Raw FRM data" tab, because the test dates associated with each data set are entered with the FRM data. When entering the data, the first measurement of the subset will be identified as "not valid" because at least 2 measurements are required for a valid data set; the "not valid" indication will normally disappear upon entering the second measurement of the data subset. Use the Comments column to enter any possibly pertinent information about any of the measurements in the data subset.
4. For each FRM data subset entered, the data are automatically checked according to the prescribed outlier test, and the mean and precision are calculated. If the data set fails to meet one or more of the prescribed validity requirements, the data subset will be identified as "not valid." Do not delete these invalid data sets on this "Raw FRM data" tab as these invalid data sets will be filtered out in other tabs. If any FRM data subset is determined to be invalid due to a known physical cause, change the "ok" in the "FRM Set Validity" column to "not valid" and enter the reason in the "Comments..." column.
5. The Raw FRM data tab allows for entry of up to 70 data sets to accommodate data from both summer and winter test seasons, if applicable, as well as some extra data sets over the minimum required, if needed. Note that the summary statistics at the bottom of the spreadsheet are for informational purposes only as they may include invalid data sets. The formal summary statistics for the valid data only are calculated correctly in other tabs.
6. Similarly, enter the candidate FEM data at the "Raw candidate data" tab for the corresponding test dates. Although the outlier tests will check the data, this outlier test is not applicable to candidate data, so the outlier test results are ignored. If the candidate data subset fails to meet one or more of the validity requirements, the data subset will be identified as "not valid." Do not delete these invalid data sets on this "Raw candidate data" tab as these invalid data sets will be filtered out in other tabs. If any candidate method data subset is determined to be invalid due to a known physical cause, change the "ok" in the "Cand. Set Validity" column to "not valid" and enter the reason in the "Comments..." column. Note that the summary statistics at the bottom of the spreadsheet are for informational purposes only as they may include invalid data sets. The formal summary statistics for the valid data sets only are calculated correctly in other tabs.
7. The mean concentrations for both FRM and candidate FEM methods for each data set are presented in the "Precision" tab, along with both the absolute precision (μg/m3) and relative precision (CV), with the later precision being the relevant one for purposes of the test results. Invalid data sets are automatically excluded from the precision calculations. If desired, the invalid data and blank cells may be hidden by using the data filter - follow the instructions in the red box in the upper right corner of the sheet.
8. The required linear regression calculations are carried out in the "Regression" tab. Invalid data sets are automatically excluded from the regression calculations. If desired, the invalid data sets and blank cells may be hidden by using the data filter - follow the instructions in the red box in the upper right of the sheet.
9. If new data are entered after filtering out the invalid data sets and blank cells, click on the ▼ in the "Data set Validity" column of both the Precision and Regression tabs, then click on "all" to include the new data. Then click on ▼ and "ok" again to filter out the invalid data sets.
10. All required data calculations for the FEM test results are presented on the "Summary" tab, along with the test requirements and a "pass" or "fail" indication. Charts present the regression results graphically. If the charts do not plot correctly, follow the instructions at the bottom of the "Summary" sheet.
Version 0.42 - 5/2/07 F. McElroy, RTI International; T. Hanley, EPA

## Title
Template for Entering Candidate Method Test
Data and Calculating Results Related to a U.S.
EPA Equivalent Method Designation Application
for PM2.5 and PM10-2.5 [40 CFR 53.35♦]
Brief Instructions: | Applicant:
● Enter applicant and candidate | Applicant name: | Example
method information here ►. | Applicant contact:
Applicant address:
● Enter test data on the "Raw | Applicant address:
FRM data" and "Raw | Applicant address:
candidate data" sheets ▼.
● Select valid data sets on the | Candidate Method:
"Precision" and "Regression" | Description: | ACME Model XYZ123 Air monitor
sheets*. | PM size category: | PM2.5
| Class (II or III): | III
● View test results on the
"Summary" sheet*.
Test site:
*Additional instructions on sheet. | Name or ID: | St. Louis test site
| Site (A, B, C, D): | C
▼More complete instructions
♦Title 40, Part 53 of the Code of Federal Regulations, revised December 18, 2006 (Federal Register, Volume 71, page 61271, October 17, 2006).
Note on G18: Please enter the FEM classification II or III.
Note on G23: Please enter the site location area: A, B, C, or D.
Note on G17: Please enter PM2.5 or PM10-2.5.
Note on G16: Enter a brief identification of the candidate method.
Note on G22: Please enter an identification for the test site.

## Raw FRM data
Reference Method (FRM) data entry
|  | Applicant: |  | Example
|  | Candidate method: |  | ACME Model XYZ123 Air monitor - PM2.5 Class III
|  | Test site: |  | St. Louis test site - (Site location C )
| Set | Test | FRM measurements, μg/m3 |  |  | Used for outlier test |  |  | Outlier tests |  |  | Number | FRM Set | Final data set |  |  | Mean | Precision | CV
| No. | Date | Sampler 1 | Sampler 2 | Sampler 3 | 1 | 2 | 3 | Sampler 1 | Sampler 2 | Sampler 3 | valid | Validity | 1 | 2 | 3 | μg/m3 | μg/m3 | % | Comments regarding FRM data set
Minimum number of sample data sets (single season) | 1 | 2007-02-03 | 43.3 | 43.9 | 44.2 | 43.3 | 43.9 | 44.2 | OK | OK | OK | 3 | ok | 43.3 | 43.9 | 44.2 | 43.8 | 0.45826 | 1.046%
| 2 | 2007-02-04 | 10.5 | 12.3 | 12 | 10.5 | 12.3 | 12 | OK | OK | OK | 3 | ok | 10.5 | 12.3 | 12 | 11.6 | 0.96437 | 8.313%
| 3 | 2007-02-05 | 36 | 36.6 | 36.7 | 36 | 36.6 | 36.7 | OK | OK | OK | 3 | ok | 36 | 36.6 | 36.7 | 36.4333 | 0.37859 | 1.039%
| 4 | 2007-02-06 | 41 | 39.6 | 41.3 | 41 | 39.6 | 41.3 | OK | OK | OK | 3 | ok | 41 | 39.6 | 41.3 | 40.6333 | 0.90738 | 2.233%
| 5 | 2007-02-07 | 17.6 | 16 | 15.7 | 17.6 | 16 | 15.7 | OK | OK | OK | 3 | ok | 17.6 | 16 | 15.7 | 16.4333 | 1.0214 | 6.216%
| 6 | 2007-02-08 | 24.8 | 25.1 | 25.7 | 24.8 | 25.1 | 25.7 | OK | OK | OK | 3 | ok | 24.8 | 25.1 | 25.7 | 25.2 | 0.45826 | 1.818%
| 7 | 2007-02-09 | 47.3 | 49.2 | 49 | 47.3 | 49.2 | 49 | OK | OK | OK | 3 | ok | 47.3 | 49.2 | 49 | 48.5 | 1.044 | 2.153%
| 8 | 2007-02-10 | 33.4 | 31.4 | 32.9 | 33.4 | 31.4 | 32.9 | OK | OK | OK | 3 | ok | 33.4 | 31.4 | 32.9 | 32.5667 | 1.0408 | 3.196%
| 9 | 2007-02-11 | 11.3 | 12.9 | 12.9 | 11.3 | 12.9 | 12.9 | OK | OK | OK | 3 | ok | 11.3 | 12.9 | 12.9 | 12.3667 | 0.92376 | 7.47%
| 10 | 2007-02-12 | 36.9 | 38.9 | 37.4 | 36.9 | 38.9 | 37.4 | OK | OK | OK | 3 | ok | 36.9 | 38.9 | 37.4 | 37.7333 | 1.0408 | 2.758%
| 11 | 2007-02-13 | 8.3 | 9.5 | 9.1 | 8.3 | 9.5 | 9.1 | OK | OK | OK | 3 | ok | 8.3 | 9.5 | 9.1 | 8.9667 | 0.61101 | 6.814%
| 12 | 2007-02-14 | 9.6 | 11.2 | 7.6 | 9.6 | 11.2 | 7.6 | OUT | OUT | OUT | 0 | NOT VALID |  |  |  |  | -- | --
| 13 | 2007-02-15 | 30.9 | 29 | 30.8 | 30.9 | 29 | 30.8 | OK | OK | OK | 3 | ok | 30.9 | 29 | 30.8 | 30.2333 | 1.0693 | 3.537%
| 14 | 2007-02-16 | 10.8 | 12.4 | 11 | 10.8 | 12.4 | 11 | OK | OK | OK | 3 | ok | 10.8 | 12.4 | 11 | 11.4 | 0.87178 | 7.647%
| 15 | 2007-02-17 | 26.5 | 25.1 | 27.3 | 26.5 | 25.1 | 27.3 | OK | OK | OK | 3 | ok | 26.5 | 25.1 | 27.3 | 26.3 | 1.1136 | 4.234%
| 16 | 2007-02-18 | 45.4 | 44.7 | 46.3 | 45.4 | 44.7 | 46.3 | OK | OK | OK | 3 | ok | 45.4 | 44.7 | 46.3 | 45.4667 | 0.80208 | 1.764%
| 17 | 2007-02-19 | 46.4 | 47.3 | 46.4 | 46.4 | 47.3 | 46.4 | OK | OK | OK | 3 | ok | 46.4 | 47.3 | 46.4 | 46.7 | 0.51962 | 1.113%
| 18 | 2007-02-20 | 11.4 | 12.7 | 9.8 | 11.4 | 12.7 | 9.8 | OK | OK | OUT | 2 | ok | 11.4 | 12.7 |  | 12.05 | 0.91924 | 7.629%
| 19 | 2007-02-21 | 40.3 | 39.9 | 40.7 | 40.3 | 39.9 | 40.7 | OK | OK | OK | 3 | ok | 40.3 | 39.9 | 40.7 | 40.3 | 0.4 | 0.9926%
| 20 | 2007-02-22 | 5.3 | 7.3 | 7 | 5.3 | 7.3 | 7 | OUT | OK | OK | 2 | ok |  | 7.3 | 7 | 7.15 | 0.21213 | 2.967%
| 21 | 2007-02-23 | 15.3 | 14.9 | 14.3 | 15.3 | 14.9 | 14.3 | OK | OK | OK | 3 | ok | 15.3 | 14.9 | 14.3 | 14.8333 | 0.50332 | 3.393%
| 22 | 2007-02-24 | 31.9 | 32.8 | 30.8 | 31.9 | 32.8 | 30.8 | OK | OK | OK | 3 | ok | 31.9 | 32.8 | 30.8 | 31.8333 | 1.0017 | 3.147%
| 23 | 2007-02-25 | 13.3 | 13.5 | 12.4 | 13.3 | 13.5 | 12.4 | OK | OK | OK | 3 | ok | 13.3 | 13.5 | 12.4 | 13.0667 | 0.58595 | 4.484%
Second season sample data sets (if required), or extra sets | 24 | 2007-02-26 | 26 | 25 | 27 | 26 | 25 | 27 | OK | OK | OK | 3 | ok | 26 | 25 | 27 | 26 | 1 | 3.846%
| 25 | 2007-02-27 | 49 | 48.2 | 47.6 | 49 | 48.2 | 47.6 | OK | OK | OK | 3 | ok | 49 | 48.2 | 47.6 | 48.2667 | 0.70238 | 1.455%
| 26 | 2007-02-28 | 46.2 | 45.1 | 45.8 | 46.2 | 45.1 | 45.8 | OK | OK | OK | 3 | ok | 46.2 | 45.1 | 45.8 | 45.7 | 0.55678 | 1.218%
| 27 | 2007-03-01 | 21.6 | 23.3 | 21 | 21.6 | 23.3 | 21 | OK | OK | OK | 3 | ok | 21.6 | 23.3 | 21 | 21.9667 | 1.193 | 5.431%
| 28 |  |  |  |  | 0 | 0 | 0
| 29 |  |  |  |  | 0 | 0 | 0
| 30 |  |  |  |  | 0 | 0 | 0
| 31 |  |  |  |  | 0 | 0 | 0
| 32 |  |  |  |  | 0 | 0 | 0
| 33 |  |  |  |  | 0 | 0 | 0
| 34 |  |  |  |  | 0 | 0 | 0
| 35 |  |  |  |  | 0 | 0 | 0
| 36 |  |  |  |  | 0 | 0 | 0
| 37 |  |  |  |  | 0 | 0 | 0
| 38 |  |  |  |  | 0 | 0 | 0
| 39 |  |  |  |  | 0 | 0 | 0
| 40 |  |  |  |  | 0 | 0 | 0
| 41 |  |  |  |  | 0 | 0 | 0
| 42 |  |  |  |  | 0 | 0 | 0
| 43 |  |  |  |  | 0 | 0 | 0
| 44 |  |  |  |  | 0 | 0 | 0
| 45 |  |  |  |  | 0 | 0 | 0
| 46 |  |  |  |  | 0 | 0 | 0
Extra sample data sets | 47 |  |  |  |  | 0 | 0 | 0
| 48 |  |  |  |  | 0 | 0 | 0
| 49 |  |  |  |  | 0 | 0 | 0
| 50 |  |  |  |  | 0 | 0 | 0
| 51 |  |  |  |  | 0 | 0 | 0
| 52 |  |  |  |  | 0 | 0 | 0
| 53 |  |  |  |  | 0 | 0 | 0
| 54 |  |  |  |  | 0 | 0 | 0
| 55 |  |  |  |  | 0 | 0 | 0
| 56 |  |  |  |  | 0 | 0 | 0
| 57 |  |  |  |  | 0 | 0 | 0
| 58 |  |  |  |  | 0 | 0 | 0
| 59 |  |  |  |  | 0 | 0 | 0
| 60 |  |  |  |  | 0 | 0 | 0
| 61 |  |  |  |  | 0 | 0 | 0
| 62 |  |  |  |  | 0 | 0 | 0
| 63 |  |  |  |  | 0 | 0 | 0
| 64 |  |  |  |  | 0 | 0 | 0
| 65 |  |  |  |  | 0 | 0 | 0
| 66 |  |  |  |  | 0 | 0 | 0
| 67 |  |  |  |  | 0 | 0 | 0
| 68 |  |  |  |  | 0 | 0 | 0
| 69 |  |  |  |  | 0 | 0 | 0
| 70 |  |  |  |  | 0 | 0 | 0
[13 empty] | Valid sets |  |  |  | All sets | Precision | CV
|  | Count | 27 | 27 | 27 | [6 empty] | Count | 26 |  |  |  | 26 | 26 | 26
|  | Max | 49 | 49.2 | 49 | [6 empty] | Max |  |  |  |  | 48.5 | 1.193 | 8.313%
|  | Min | 5.3 | 7.3 | 7 | [6 empty] | Min |  |  |  |  | 7.15 | 0.21213 | 0.9926%
|  | Mean | 27.4185 | 27.6963 | 27.5074 | [6 empty] | Mean |  |  |  |  | 28.2885 | 0.78075 | 3.689%
Note on U9: Precision of the FRM data set, RPj [§53.35(e)(1), Equation 13]. CV = std. dev.(Col. T)/mean(Col. S). A value in red indicates the data set is not valid for test purposes.
Note on T9: Precision as standard deviation of the FRM data set. (A value in red indicates the data set is not valid for test purposes.)
Note on D11: Enter date applicable to the test data set.
Note on S82: Unofficial - includes non-valid data sets. See the "Precision" or "Regression" sheets for official test purposes.
Note on K9: Tests for FRM outliers [§53.35(d)(1)].
Note on N9: The number of valid FRM measurements in the set, based on the results of the outlier tests [§53.35(d)(1)(iv)] (but not considering the mean FRM concentration level).
Note on O9: Determines the validity of the FRM data set based on having at least 2 valid FRM measurements [§53.35(c)(2)] within the mean acceptable concentration range of 3 - 200 μg/m3 [Table C-4].
Note on S9: Mean of the valid FRM measurements, Rj [§53.35(d)(2), Equation 11]. A value in red indicates the data set is not valid for test purposes.
Note on R9: Valid FRM measurements, based on the results of the outlier test.
Note on P9: Valid FRM measurements, based on the results of the outlier test.
Note on Q9: Valid FRM measurements, based on the results of the outlier test.
Note on T82: Unofficial - includes non-valid data sets.
Note on U82: Unofficial - includes non-valid data sets. See the "Precision" sheet for precision for test purposes.
Note on D9: Enter date applicable to the test data set.
Note on U83: Unofficial - includes non-valid data sets. See the "Precision" sheet for precision for test purposes.
Note on U84: Unofficial - includes non-valid data sets. See the "Precision" sheet for precision for test purposes.
Note on U85: Unofficial - includes non-valid data sets. See the "Precision" sheet for precision for test purposes.
Note on U86: Unofficial - includes non-valid data sets. See the "Precision" sheet for precision for test purposes.
Note on T83: Unofficial - includes non-valid data sets.
Note on T84: Unofficial - includes non-valid data sets.
Note on T85: Unofficial - includes non-valid data sets.
Note on T86: Unofficial - includes non-valid data sets.
Note on S83: Unofficial - includes non-valid data sets. See the "Precision" or "Regression" sheets for official test purposes.
Note on S84: Unofficial - includes non-valid data sets. See the "Precision" or "Regression" sheets for official test purposes.
Note on S85: Unofficial - includes non-valid data sets. See the "Precision" or "Regression" sheets for official test purposes.
Note on S86: Unofficial - includes non-valid data sets. See the "Precision" or "Regression" sheets for official test purposes.
Note on O82: See the "Precision" or "Regression" sheets for official calculations for valid data sets.
Note on L9: Tests for FRM outliers [§53.35(d)(1)].
Note on M9: Tests for FRM outliers [§53.35(d)(1)].

## Raw candidate data
Candidate Method (FEM) data entry
|  | Applicant: |  | Example
|  | Candidate method: |  | ACME Model XYZ123 Air monitor - PM2.5 Class III
|  | Test site: |  | St. Louis test site - (Site location C )
| Set | Test | Candidate measurements, μg/m3 |  |  | Used for outlier test |  |  | Outlier tests: NOT APPLICABLE |  |  | Values | Cand. Set | Mean | Precision | CV
| No. | Date | Unit 1 | Unit 2 | Unit 3 | 1 | 2 | 3 | Sampler 1 | Sampler 2 | Sampler 3 | Available | Validity | μg/m3 | μg/m3 | % | Comments regarding candidate data set
Minimum number of sample data sets (single season) | 1 | 2007-02-03 | 44 | 42.5 | 42.9 | 44 | 42.5 | 42.9 | OK | OK | OK | 3 | ok | 43.1333 | 0.77675 | 1.801%
| 2 | 2007-02-04 | 12.2 | 12.4 | 8.9 | 12.2 | 12.4 | 8.9 | OK | OK | OUT | 3 | ok | 11.1667 | 1.9655 | 17.6%
| 3 | 2007-02-05 | 38.1 | 36.5 | 36.7 | 38.1 | 36.5 | 36.7 | OK | OK | OK | 3 | ok | 37.1 | 0.87178 | 2.35%
| 4 | 2007-02-06 | 41.8 | 44.2 | 40 | 41.8 | 44.2 | 40 | OK | OK | OK | 3 | ok | 42 | 2.1071 | 5.017%
| 5 | 2007-02-07 | 20.4 | 18.1 | 20.1 | 20.4 | 18.1 | 20.1 | OK | OK | OK | 3 | ok | 19.5333 | 1.2503 | 6.401%
| 6 | 2007-02-08 | 22.8 | 27 | 22.8 | 22.8 | 27 | 22.8 | OK | OUT | OK | 3 | ok | 24.2 | 2.4249 | 10.02%
| 7 | 2007-02-09 | 45.4 | 48.7 | 45.6 | 45.4 | 48.7 | 45.6 | OK | OK | OK | 3 | ok | 46.5667 | 1.8502 | 3.973%
| 8 | 2007-02-10 | 33.7 | 35.6 | 34.6 | 33.7 | 35.6 | 34.6 | OK | OK | OK | 3 | ok | 34.6333 | 0.95044 | 2.744%
| 9 | 2007-02-11 | 11.1 | 9 | 14.2 | 11.1 | 9 | 14.2 | OUT | OUT | OUT | 3 | ok | 11.4333 | 2.616 | 22.88%
| 10 | 2007-02-12 | 36.7 | 35.9 | 37.7 | 36.7 | 35.9 | 37.7 | OK | OK | OK | 3 | ok | 36.7667 | 0.90185 | 2.453%
| 11 | 2007-02-13 | 7.8 | 9.7 | 6.2 | 7.8 | 9.7 | 6.2 | OUT | OUT | OUT | 3 | ok | 7.9 | 1.7521 | 22.18%
| 12 | 2007-02-14 | 10.5 | 10.2 | 13.1 | 10.5 | 10.2 | 13.1 | OK | OK | OUT | 3 | NOT VALID | 11.2667 | 1.5948 | 14.15%
| 13 | 2007-02-15 | 30 | 33 | 29.9 | 30 | 33 | 29.9 | OK | OK | OK | 3 | ok | 30.9667 | 1.7616 | 5.689%
| 14 | 2007-02-16 | 13.2 | 12.5 | 10.9 | 13.2 | 12.5 | 10.9 | OK | OK | OK | 3 | ok | 12.2 | 1.179 | 9.664%
| 15 | 2007-02-17 | 26.1 | 24.5 | 29.5 | 26.1 | 24.5 | 29.5 | OK | OK | OK | 3 | ok | 26.7 | 2.5534 | 9.563%
| 16 | 2007-02-18 | 46.4 | 46.1 | 47.8 | 46.4 | 46.1 | 47.8 | OK | OK | OK | 3 | ok | 46.7667 | 0.90738 | 1.94%
| 17 | 2007-02-19 | 45.8 | 49.7 | 45.5 | 45.8 | 49.7 | 45.5 | OK | OK | OK | 3 | ok | 47 | 2.3431 | 4.985%
| 18 | 2007-02-20 | 12.1 | 11.2 | 11.8 | 12.1 | 11.2 | 11.8 | OK | OK | OK | 3 | ok | 11.7 | 0.45826 | 3.917%
| 19 | 2007-02-21 | 42.6 | 38.2 | 40.4 | 42.6 | 38.2 | 40.4 | OK | OK | OK | 3 | ok | 40.4 | 2.2 | 5.446%
| 20 | 2007-02-22 | 6.9 | 3.2 | 7.3 | 6.9 | 3.2 | 7.3 | OK | OUT | OK | 3 | ok | 5.8 | 2.2605 | 38.97%
| 21 | 2007-02-23 | 13.4 | 16.7 | 16.1 | 13.4 | 16.7 | 16.1 | OUT | OK | OK | 3 | ok | 15.4 | 1.7578 | 11.41%
| 22 | 2007-02-24 | 29.8 | 30.6 | 31.3 | 29.8 | 30.6 | 31.3 | OK | OK | OK | 3 | ok | 30.5667 | 0.75056 | 2.455%
| 23 | 2007-02-25 | 16.3 | 14.3 | 11.6 | 16.3 | 14.3 | 11.6 | OK | OK | OUT | 3 | ok | 14.0667 | 2.3587 | 16.77%
Second season sample data sets (if required), or extra sets | 24 | 2007-02-26 | 25.1 | 29.1 | 27.1 | 25.1 | 29.1 | 27.1 | OK | OK | OK | 3 | ok | 27.1 | 2 | 7.38%
| 25 | 2007-02-27 | 48 | 47 | 51.5 | 48 | 47 | 51.5 | OK | OK | OK | 3 | ok | 48.8333 | 2.3629 | 4.839%
| 26 | 2007-02-28 | 46.3 | 48.8 | 47.4 | 46.3 | 48.8 | 47.4 | OK | OK | OK | 3 | ok | 47.5 | 1.253 | 2.638%
| 27 | 2007-03-01 | 20.3 | 21.8 | 23.2 | 20.3 | 21.8 | 23.2 | OK | OK | OK | 3 | ok | 21.7667 | 1.4503 | 6.663%
| 28 |  | Used for outlier test 1: 0; Used for outlier test 2: 0; Used for outlier test 3: 0
| 29 |  | Used for outlier test 1: 0; Used for outlier test 2: 0; Used for outlier test 3: 0
| 30 |  | Used for outlier test 1: 0; Used for outlier test 2: 0; Used for outlier test 3: 0
| 31 |  | Used for outlier test 1: 0; Used for outlier test 2: 0; Used for outlier test 3: 0
| 32 |  | Used for outlier test 1: 0; Used for outlier test 2: 0; Used for outlier test 3: 0
| 33 |  | Used for outlier test 1: 0; Used for outlier test 2: 0; Used for outlier test 3: 0
| 34 |  | Used for outlier test 1: 0; Used for outlier test 2: 0; Used for outlier test 3: 0
| 35 |  | Used for outlier test 1: 0; Used for outlier test 2: 0; Used for outlier test 3: 0
| 36 |  | Used for outlier test 1: 0; Used for outlier test 2: 0; Used for outlier test 3: 0
| 37 |  | Used for outlier test 1: 0; Used for outlier test 2: 0; Used for outlier test 3: 0
| 38 |  | Used for outlier test 1: 0; Used for outlier test 2: 0; Used for outlier test 3: 0
| 39 |  | Used for outlier test 1: 0; Used for outlier test 2: 0; Used for outlier test 3: 0
| 40 |  | Used for outlier test 1: 0; Used for outlier test 2: 0; Used for outlier test 3: 0
| 41 |  | Used for outlier test 1: 0; Used for outlier test 2: 0; Used for outlier test 3: 0
| 42 |  | Used for outlier test 1: 0; Used for outlier test 2: 0; Used for outlier test 3: 0
| 43 |  | Used for outlier test 1: 0; Used for outlier test 2: 0; Used for outlier test 3: 0
| 44 |  | Used for outlier test 1: 0; Used for outlier test 2: 0; Used for outlier test 3: 0
| 45 |  | Used for outlier test 1: 0; Used for outlier test 2: 0; Used for outlier test 3: 0
| 46 |  | Used for outlier test 1: 0; Used for outlier test 2: 0; Used for outlier test 3: 0
Extra sample data sets | 47 |  | Used for outlier test 1: 0; Used for outlier test 2: 0; Used for outlier test 3: 0
| 48 |  | Used for outlier test 1: 0; Used for outlier test 2: 0; Used for outlier test 3: 0
| 49 |  | Used for outlier test 1: 0; Used for outlier test 2: 0; Used for outlier test 3: 0
| 50 |  | Used for outlier test 1: 0; Used for outlier test 2: 0; Used for outlier test 3: 0
| 51 |  | Used for outlier test 1: 0; Used for outlier test 2: 0; Used for outlier test 3: 0
| 52 |  | Used for outlier test 1: 0; Used for outlier test 2: 0; Used for outlier test 3: 0
| 53 |  | Used for outlier test 1: 0; Used for outlier test 2: 0; Used for outlier test 3: 0
| 54 |  | Used for outlier test 1: 0; Used for outlier test 2: 0; Used for outlier test 3: 0
| 55 |  | Used for outlier test 1: 0; Used for outlier test 2: 0; Used for outlier test 3: 0
| 56 |  | Used for outlier test 1: 0; Used for outlier test 2: 0; Used for outlier test 3: 0
| 57 |  | Used for outlier test 1: 0; Used for outlier test 2: 0; Used for outlier test 3: 0
| 58 |  | Used for outlier test 1: 0; Used for outlier test 2: 0; Used for outlier test 3: 0
| 59 |  | Used for outlier test 1: 0; Used for outlier test 2: 0; Used for outlier test 3: 0
| 60 |  | Used for outlier test 1: 0; Used for outlier test 2: 0; Used for outlier test 3: 0
| 61 |  | Used for outlier test 1: 0; Used for outlier test 2: 0; Used for outlier test 3: 0
| 62 |  | Used for outlier test 1: 0; Used for outlier test 2: 0; Used for outlier test 3: 0
| 63 |  | Used for outlier test 1: 0; Used for outlier test 2: 0; Used for outlier test 3: 0
| 64 |  | Used for outlier test 1: 0; Used for outlier test 2: 0; Used for outlier test 3: 0
| 65 |  | Used for outlier test 1: 0; Used for outlier test 2: 0; Used for outlier test 3: 0
| 66 |  | Used for outlier test 1: 0; Used for outlier test 2: 0; Used for outlier test 3: 0
| 67 |  | Used for outlier test 1: 0; Used for outlier test 2: 0; Used for outlier test 3: 0
| 68 |  | Used for outlier test 1: 0; Used for outlier test 2: 0; Used for outlier test 3: 0
| 69 |  | Used for outlier test 1: 0; Used for outlier test 2: 0; Used for outlier test 3: 0
| 70 |  | Used for outlier test 1: 0; Used for outlier test 2: 0; Used for outlier test 3: 0
[13 empty] | Valid sets | All sets | Precision | CV
|  | Count | Candidate measurements, μg/m3 Unit 1: 27; Candidate measurements, μg/m3 Unit 2: 27; Candidate measurements, μg/m3 Unit 3: 27 |  |  |  | Count | 26 | 27 | 27 | 27
|  | Max | Candidate measurements, μg/m3 Unit 1: 48; Candidate measurements, μg/m3 Unit 2: 49.7; Candidate measurements, μg/m3 Unit 3: 51.5 |  |  |  | Max |  | 48.8333 | 2.616 | 38.97%
|  | Min | Candidate measurements, μg/m3 Unit 1: 6.9; Candidate measurements, μg/m3 Unit 2: 3.2; Candidate measurements, μg/m3 Unit 3: 6.2 |  |  |  | Min |  | 5.8 | 0.45826 | 1.801%
|  | Mean | Candidate measurements, μg/m3 Unit 1: 27.6593; Candidate measurements, μg/m3 Unit 2: 28.0185; Candidate measurements, μg/m3 Unit 3: 27.9296 |  |  |  | Mean |  | 27.8691 | 1.654 | 9.034%
Note on Q9: Precision as the standard deviation of the candidate method measurements.
Note on R9: Precision of the candidate method data set, CPj [§53.35(f)(1), Equation 15]. CV = std. dev.(Col. Q)/mean(Col. P).
Note on D11: Enter dates on the Raw FRM data sheet.
Note on M9: This information is provided only for information purposes; it is not used to determine validity of candidate method measurements.
Note on D12: Enter dates on the Raw FRM data sheet.
Note on D13: Enter dates on the Raw FRM data sheet.
Note on D14: Enter dates on the Raw FRM data sheet.
Note on D15: Enter dates on the Raw FRM data sheet.
Note on D16: Enter dates on the Raw FRM data sheet.
Note on D17: Enter dates on the Raw FRM data sheet.
Note on D18: Enter dates on the Raw FRM data sheet.
Note on D19: Enter dates on the Raw FRM data sheet.
Note on D20: Enter dates on the Raw FRM data sheet.
Note on D21: Enter dates on the Raw FRM data sheet.
Note on D22: Enter dates on the Raw FRM data sheet.
Note on D23: Enter dates on the Raw FRM data sheet.
Note on D24: Enter dates on the Raw FRM data sheet.
Note on D25: Enter dates on the Raw FRM data sheet.
Note on D26: Enter dates on the Raw FRM data sheet.
Note on D27: Enter dates on the Raw FRM data sheet.
Note on D28: Enter dates on the Raw FRM data sheet.
Note on D29: Enter dates on the Raw FRM data sheet.
Note on D30: Enter dates on the Raw FRM data sheet.
Note on D31: Enter dates on the Raw FRM data sheet.
Note on D32: Enter dates on the Raw FRM data sheet.
Note on D33: Enter dates on the Raw FRM data sheet.
Note on D34: Enter dates on the Raw FRM data sheet.
Note on D35: Enter dates on the Raw FRM data sheet.
Note on D36: Enter dates on the Raw FRM data sheet.
Note on D37: Enter dates on the Raw FRM data sheet.
Note on D38: Enter dates on the Raw FRM data sheet.
Note on D39: Enter dates on the Raw FRM data sheet.
Note on D40: Enter dates on the Raw FRM data sheet.
Note on D41: Enter dates on the Raw FRM data sheet.
Note on D42: Enter dates on the Raw FRM data sheet.
Note on D43: Enter dates on the Raw FRM data sheet.
Note on D44: Enter dates on the Raw FRM data sheet.
Note on D45: Enter dates on the Raw FRM data sheet.
Note on D46: Enter dates on the Raw FRM data sheet.
Note on D47: Enter dates on the Raw FRM data sheet.
Note on D48: Enter dates on the Raw FRM data sheet.
Note on D49: Enter dates on the Raw FRM data sheet.
Note on D50: Enter dates on the Raw FRM data sheet.
Note on D51: Enter dates on the Raw FRM data sheet.
Note on D52: Enter dates on the Raw FRM data sheet.
Note on D53: Enter dates on the Raw FRM data sheet.
Note on D54: Enter dates on the Raw FRM data sheet.
Note on D55: Enter dates on the Raw FRM data sheet.
Note on D56: Enter dates on the Raw FRM data sheet.
Note on D57: Enter dates on the Raw FRM data sheet.
Note on O9: Validity of the candidate data set is based on having at least 2 valid candidate measurements [§53.35(c)(2)] and a valid FRM data set with mean FRM concentration measurement between 3 and 300 μm/m3 [Table C-4].
Note on N9: The outlier test does not apply to candidate method measurements.
Note on P9: Mean of the candidate method conentration measurements, Cj [§53.35(d)(4), Equation 12].
Note on O82: See the "Precision" or "Regression" sheets for official calculations for valid data sets.
Note on P82: Unofficial - includes non-valid data sets. See the "Precision" or "Regression" sheets for official test purposes.
Note on P83: Unofficial - includes non-valid data sets. See the "Precision" or "Regression" sheets for official test purposes.
Note on P84: Unofficial - includes non-valid data sets. See the "Precision" or "Regression" sheets for official test purposes.
Note on P85: Unofficial - includes non-valid data sets. See the "Precision" or "Regression" sheets for official test purposes.
Note on P86: Unofficial - includes non-valid data sets. See the "Precision" or "Regression" sheets for official test purposes.
Note on Q82: Unofficial - includes non-valid data sets.
Note on Q83: Unofficial - includes non-valid data sets.
Note on Q84: Unofficial - includes non-valid data sets.
Note on Q85: Unofficial - includes non-valid data sets.
Note on Q86: Unofficial - includes non-valid data sets.
Note on R82: Unofficial - includes non-valid data sets. See the "Precision" sheet for precision for test purposes.
Note on R83: Unofficial - includes non-valid data sets. See the "Precision" sheet for precision for test purposes.
Note on R84: Unofficial - includes non-valid data sets. See the "Precision" sheet for precision for test purposes.
Note on R85: Unofficial - includes non-valid data sets. See the "Precision" sheet for precision for test purposes.
Note on R86: Unofficial - includes non-valid data sets. See the "Precision" sheet for precision for test purposes.
Note on D58: Enter dates on the Raw FRM data sheet.
Note on D59: Enter dates on the Raw FRM data sheet.
Note on D60: Enter dates on the Raw FRM data sheet.
Note on D61: Enter dates on the Raw FRM data sheet.
Note on D62: Enter dates on the Raw FRM data sheet.
Note on D63: Enter dates on the Raw FRM data sheet.
Note on D64: Enter dates on the Raw FRM data sheet.
Note on D65: Enter dates on the Raw FRM data sheet.
Note on D66: Enter dates on the Raw FRM data sheet.
Note on D67: Enter dates on the Raw FRM data sheet.
Note on D68: Enter dates on the Raw FRM data sheet.
Note on D69: Enter dates on the Raw FRM data sheet.
Note on D70: Enter dates on the Raw FRM data sheet.
Note on D71: Enter dates on the Raw FRM data sheet.
Note on D72: Enter dates on the Raw FRM data sheet.
Note on D73: Enter dates on the Raw FRM data sheet.
Note on D74: Enter dates on the Raw FRM data sheet.
Note on D75: Enter dates on the Raw FRM data sheet.
Note on D76: Enter dates on the Raw FRM data sheet.
Note on D77: Enter dates on the Raw FRM data sheet.
Note on D78: Enter dates on the Raw FRM data sheet.
Note on D79: Enter dates on the Raw FRM data sheet.
Note on D80: Enter dates on the Raw FRM data sheet.

## Precision
Calculation of Precision [53.35(e) and (f)]
|  | Applicant: |  |  |  |  | Example | [6 empty] | Click on the ▼ arrow
|  | Candidate method: |  |  |  |  | ACME Model XYZ123 Air monitor - PM2.5 Class III | [6 empty] | in the "Validity" column,
|  | Test site: |  |  |  |  | St. Louis test site - (Site location C ) | [6 empty] | then click on "ok" to select
only valid data sets.
| Set | Test | # sets |  | Data set |  | Data set mean, μg/m3 |  | Data set precision, μg/m3 |  | Relative precision (CV) |  |  | (Click "all" to restore all
| No. | Date | FRM | Cand | Validity | FRM | FRM | Candidate | FRM | Candidate | FRM | Candidate | Candidate | data sets.)
Minimum number of sample data sets (single season) | 1 | 2007-02-03 | 3 | 3 | ok | 43.8 | 43.8 | 43.1333 | 0.45826 | 0.77675 | 1.046% | 1.801% | 1.801%
| 2 | 2007-02-04 | 3 | 3 | ok | 11.6 | 11.6 | 11.1667 | 0.96437 | 1.9655 | 8.313% | 17.6% | 17.6%
| 3 | 2007-02-05 | 3 | 3 | ok | 36.4333 | 36.4333 | 37.1 | 0.37859 | 0.87178 | 1.039% | 2.35% | 2.35%
| 4 | 2007-02-06 | 3 | 3 | ok | 40.6333 | 40.6333 | 42 | 0.90738 | 2.1071 | 2.233% | 5.017% | 5.017%
| 5 | 2007-02-07 | 3 | 3 | ok | 16.4333 | 16.4333 | 19.5333 | 1.0214 | 1.2503 | 6.216% | 6.401% | 6.401%
| 6 | 2007-02-08 | 3 | 3 | ok | 25.2 | 25.2 | 24.2 | 0.45826 | 2.4249 | 1.818% | 10.02% | 10.02%
| 7 | 2007-02-09 | 3 | 3 | ok | 48.5 | 48.5 | 46.5667 | 1.044 | 1.8502 | 2.153% | 3.973% | 3.973%
| 8 | 2007-02-10 | 3 | 3 | ok | 32.5667 | 32.5667 | 34.6333 | 1.0408 | 0.95044 | 3.196% | 2.744% | 2.744%
| 9 | 2007-02-11 | 3 | 3 | ok | 12.3667 | 12.3667 | 11.4333 | 0.92376 | 2.616 | 7.47% | 22.88% | 22.88%
| 10 | 2007-02-12 | 3 | 3 | ok | 37.7333 | 37.7333 | 36.7667 | 1.0408 | 0.90185 | 2.758% | 2.453% | 2.453%
| 11 | 2007-02-13 | 3 | 3 | ok | 8.9667 | 8.9667 | 7.9 | 0.61101 | 1.7521 | 6.814% | 22.18% | 22.18%
| 12 | 2007-02-14 |  |  | NOT VALID | FALSE | [5 empty] | FALSE
| 13 | 2007-02-15 | 3 | 3 | ok | 30.2333 | 30.2333 | 30.9667 | 1.0693 | 1.7616 | 3.537% | 5.689% | 5.689%
| 14 | 2007-02-16 | 3 | 3 | ok | 11.4 | 11.4 | 12.2 | 0.87178 | 1.179 | 7.647% | 9.664% | 9.664%
| 15 | 2007-02-17 | 3 | 3 | ok | 26.3 | 26.3 | 26.7 | 1.1136 | 2.5534 | 4.234% | 9.563% | 9.563%
| 16 | 2007-02-18 | 3 | 3 | ok | 45.4667 | 45.4667 | 46.7667 | 0.80208 | 0.90738 | 1.764% | 1.94% | 1.94%
| 17 | 2007-02-19 | 3 | 3 | ok | 46.7 | 46.7 | 47 | 0.51962 | 2.3431 | 1.113% | 4.985% | 4.985%
| 18 | 2007-02-20 | 2 | 3 | ok | 12.05 | 12.05 | 11.7 | 0.91924 | 0.45826 | 7.629% | 3.917% | 3.917%
| 19 | 2007-02-21 | 3 | 3 | ok | 40.3 | 40.3 | 40.4 | 0.4 | 2.2 | 0.9926% | 5.446% | 5.446%
| 20 | 2007-02-22 | 2 | 3 | ok | 7.15 | 7.15 | 5.8 | 0.21213 | 2.2605 | 2.967% | 38.97% | 38.97%
| 21 | 2007-02-23 | 3 | 3 | ok | 14.8333 | 14.8333 | 15.4 | 0.50332 | 1.7578 | 3.393% | 11.41% | 11.41%
| 22 | 2007-02-24 | 3 | 3 | ok | 31.8333 | 31.8333 | 30.5667 | 1.0017 | 0.75056 | 3.147% | 2.455% | 2.455%
| 23 | 2007-02-25 | 3 | 3 | ok | 13.0667 | 13.0667 | 14.0667 | 0.58595 | 2.3587 | 4.484% | 16.77% | 16.77%
Second season sample data sets (if required), or extra sets | 24 | 2007-02-26 | 3 | 3 | ok | 26 | 26 | 27.1 | 1 | 2 | 3.846% | 7.38% | 7.38%
| 25 | 2007-02-27 | 3 | 3 | ok | 48.2667 | 48.2667 | 48.8333 | 0.70238 | 2.3629 | 1.455% | 4.839% | 4.839%
| 26 | 2007-02-28 | 3 | 3 | ok | 45.7 | 45.7 | 47.5 | 0.55678 | 1.253 | 1.218% | 2.638% | 2.638%
| 27 | 2007-03-01 | 3 | 3 | ok | 21.9667 | 21.9667 | 21.7667 | 1.193 | 1.4503 | 5.431% | 6.663% | 6.663%
| 28 |  |  |  |  | FALSE | [5 empty] | FALSE
| 29 |  |  |  |  | FALSE | [5 empty] | FALSE
| 30 |  |  |  |  | FALSE | [5 empty] | FALSE
| 31 |  |  |  |  | FALSE | [5 empty] | FALSE
| 32 |  |  |  |  | FALSE | [5 empty] | FALSE
| 33 |  |  |  |  | FALSE | [5 empty] | FALSE
| 34 |  |  |  |  | FALSE | [5 empty] | FALSE
| 35 |  |  |  |  | FALSE | [5 empty] | FALSE
| 36 |  |  |  |  | FALSE | [5 empty] | FALSE
| 37 |  |  |  |  | FALSE | [5 empty] | FALSE
| 38 |  |  |  |  | FALSE | [5 empty] | FALSE
| 39 |  |  |  |  | FALSE | [5 empty] | FALSE
| 40 |  |  |  |  | FALSE | [5 empty] | FALSE
| 41 |  |  |  |  | FALSE | [5 empty] | FALSE
| 42 |  |  |  |  | FALSE | [5 empty] | FALSE
| 43 |  |  |  |  | FALSE | [5 empty] | FALSE
| 44 |  |  |  |  | FALSE | [5 empty] | FALSE
| 45 |  |  |  |  | FALSE | [5 empty] | FALSE
| 46 |  |  |  |  | FALSE | [5 empty] | FALSE
Extra sets | 47 |  |  |  |  | FALSE | [5 empty] | FALSE
| 48 |  |  |  |  | FALSE | [5 empty] | FALSE
| 49 |  |  |  |  | FALSE | [5 empty] | FALSE
| 50 |  |  |  |  | FALSE | [5 empty] | FALSE
| 51 |  |  |  |  | FALSE | [5 empty] | FALSE
| 52 |  |  |  |  | FALSE | [5 empty] | FALSE
| 53 |  |  |  |  | FALSE | [5 empty] | FALSE
| 54 |  |  |  |  | FALSE | [5 empty] | FALSE
| 55 |  |  |  |  | FALSE | [5 empty] | FALSE
| 56 |  |  |  |  | FALSE | [5 empty] | FALSE
| 57 |  |  |  |  | FALSE | [5 empty] | FALSE
| 58 |  |  |  |  | FALSE | [5 empty] | FALSE
| 59 |  |  |  |  | FALSE | [5 empty] | FALSE
| 60 |  |  |  |  | FALSE | [5 empty] | FALSE
| 61 |  |  |  |  | FALSE | [5 empty] | FALSE
| 62 |  |  |  |  | FALSE | [5 empty] | FALSE
| 63 |  |  |  |  | FALSE | [5 empty] | FALSE
| 64 |  |  |  |  | FALSE | [5 empty] | FALSE
| 65 |  |  |  |  | FALSE | [5 empty] | FALSE
| 66 |  |  |  |  | FALSE | [5 empty] | FALSE
| 67 |  |  |  |  | FALSE | [5 empty] | FALSE
| 68 |  |  |  |  | FALSE | [5 empty] | FALSE
| 69 |  |  |  |  | FALSE | [5 empty] | FALSE
| 70 |  |  |  |  | FALSE | [5 empty] | FALSE
[7 empty] | Data set mean, μg/m3 |  | Data set precision, μg/m3 |  | Relative precision (CV)
[7 empty] | FRM | Candidate | FRM | Candidate | FRM |  | Candidate
[5 empty] | No. valid data sets: |  | 26 | 26 | 26 | 26 | 26 |  | 26
[5 empty] | Mean: |  | 28.2885 | 28.5077 | 0.78075 | 1.6563 | 3.689% |  | 8.837%
[5 empty] | Maximum: |  | 48.5 | 48.8333 | 1.193 | 2.616 | 8.313% |  | 38.97%
[5 empty] | Minimum: |  | 7.15 | 5.8 | 0.21213 | 0.45826 | 0.9926% |  | 1.801%
[5 empty] | Candidate/FRM Ratio: |  |  | 100.8% |  | 212.1% |  |  | 239.5%
[11 empty] | FRM |  | Candidate
[7 empty] | RMS Precision for site |  |  |  | 4.357% |  | 12.23%
[7 empty] | Test requirements - PM2.5 Class III |  |  |  | 10% |  | 15%
[7 empty] | Precision Test Results for site |  |  |  | OK |  | PASS
26
Number of valid data sets available:
[7 empty] | Number of valid data sets | [5 empty] | 23
required for Class III, location C:
[7 empty] | The number of |  |  |  | OK
valid data sets for this test is:
Additional data sets needed:
Note on M89: FRM precision for the test site is the root mean square (RMS) of the precisions of the FRM data sets [§53.35(e)(2), Equation 14].
Note on O89: Candidate method precision for the test site is the root mean square (RMS) of the precisions of the candidate data sets [§53.35(f)(2), Equation 16].
Note on G7: Validity requires that both the corresponding FRM and candidate data sets are valid [§53.35(c)(2)].
Note on J7: Means of the FRM and Candidate method measurements for valid data sets [§53.35(d)(2) and (4), Equations 11 and 12, respectively].
Note on L7: Precisions as the standard deviation of the FRM and candidate method measurements, respectively. These precision values are not used directly for test purposes.
Note on O7: Precisions of the FRM and Candidate method valid data sets, RPj and CPj [§53.35(e)(1) and (f)(1), Equations13 and 15]. CV = std. dev.(Col. J or K)/mean(Col. H or I).
Note on J80: Means of the FRM and Candidate method measurements for valid data sets [§53.35(d)(2) and (4), Equations 11 and 12, respectively].
Note on L80: Precisions as the standard deviation of the FRM and candidate method measurements, respectively. These precision values are not used directly for test purposes.
Note on O80: Precisions of the FRM and Candidate method valid data sets, RPj and CPj [§53.35(e)(1) and (f)(1), Equations13 and 15]. CV = std. dev.(Col. J or K)/mean(Col. H or I).
Note on I90: Requirement for method precision [Table C-4].
Note on M90: Requirement for FRM method precision [Table C-4].
Note on O90: Requirement for Candidate method precision [Table C-4].
Note on L89: FRM precision for the test site is the root mean square (RMS) of the precisions of the FRM or Candidate data sets [§53.35(e)(2) or (f)(2), Equations 14 or 16].
Note on L95: At least 46 valid data sets required for site location A, 23 for site locations B, C, and D [§53.35(b)(2) and (c)(2), and Table C-5]
Note on O95: At least 46 valid data sets required for site location A, 23 for site locations B, C, and D [§53.35(b)(2) and (c)(2), and Table C-5]
Note on L91: Test results are based only on the calculated precision, not considering whether the number of valid data sets is sufficient.
Note on M91: Test result is based only on the calculated precision, not considering whether the number of valid data sets is sufficient.
Note on O91: Test result is based only on the calculated precision, not considering whether the number of valid data sets is sufficient.

## Regression
Calculation of Slope, Intercept, and Correlation [53.35(g) and (h)]
|  | Applicant: |  |  |  |  | Example
|  | Candidate method: |  |  |  |  | ACME Model XYZ123 Air monitor - PM2.5 Class III
|  | Test site: |  |  |  |  | St. Louis test site - (Site location C )
| Set | Test | # sets |  | Data set |  | Data set mean, μg/m3 |  |  | Please click on the ▼ arrow in the "Validity" column, then click on
| No. | Date | FRM | Cand | Validity | FRM | FRM | Candidate | Candidate | "ok" to select only valid data sets. (Click "all" to restore all sets.)
Minimum number of sample data sets (single season) | 1 | 2007-02-03 | 3 | 3 | ok | 43.8 | 43.8 | 43.1333 | 43.1333
| 2 | 2007-02-04 | 3 | 3 | ok | 11.6 | 11.6 | 11.1667 | 11.1667
| 3 | 2007-02-05 | 3 | 3 | ok | 36.4333 | 36.4333 | 37.1 | 37.1
| 4 | 2007-02-06 | 3 | 3 | ok | 40.6333 | 40.6333 | 42 | 42
| 5 | 2007-02-07 | 3 | 3 | ok | 16.4333 | 16.4333 | 19.5333 | 19.5333
| 6 | 2007-02-08 | 3 | 3 | ok | 25.2 | 25.2 | 24.2 | 24.2
| 7 | 2007-02-09 | 3 | 3 | ok | 48.5 | 48.5 | 46.5667 | 46.5667
| 8 | 2007-02-10 | 3 | 3 | ok | 32.5667 | 32.5667 | 34.6333 | 34.6333
| 9 | 2007-02-11 | 3 | 3 | ok | 12.3667 | 12.3667 | 11.4333 | 11.4333
| 10 | 2007-02-12 | 3 | 3 | ok | 37.7333 | 37.7333 | 36.7667 | 36.7667
| 11 | 2007-02-13 | 3 | 3 | ok | 8.9667 | 8.9667 | 7.9 | 7.9
| 12 | 2007-02-14 |  |  | NOT VALID | FALSE |  | FALSE
| 13 | 2007-02-15 | 3 | 3 | ok | 30.2333 | 30.2333 | 30.9667 | 30.9667
| 14 | 2007-02-16 | 3 | 3 | ok | 11.4 | 11.4 | 12.2 | 12.2
| 15 | 2007-02-17 | 3 | 3 | ok | 26.3 | 26.3 | 26.7 | 26.7
| 16 | 2007-02-18 | 3 | 3 | ok | 45.4667 | 45.4667 | 46.7667 | 46.7667
| 17 | 2007-02-19 | 3 | 3 | ok | 46.7 | 46.7 | 47 | 47
| 18 | 2007-02-20 | 2 | 3 | ok | 12.05 | 12.05 | 11.7 | 11.7
| 19 | 2007-02-21 | 3 | 3 | ok | 40.3 | 40.3 | 40.4 | 40.4
| 20 | 2007-02-22 | 2 | 3 | ok | 7.15 | 7.15 | 5.8 | 5.8
| 21 | 2007-02-23 | 3 | 3 | ok | 14.8333 | 14.8333 | 15.4 | 15.4
| 22 | 2007-02-24 | 3 | 3 | ok | 31.8333 | 31.8333 | 30.5667 | 30.5667
| 23 | 2007-02-25 | 3 | 3 | ok | 13.0667 | 13.0667 | 14.0667 | 14.0667
Second season sample data sets (if required), or extra sets | 24 | 2007-02-26 | 3 | 3 | ok | 26 | 26 | 27.1 | 27.1
| 25 | 2007-02-27 | 3 | 3 | ok | 48.2667 | 48.2667 | 48.8333 | 48.8333
| 26 | 2007-02-28 | 3 | 3 | ok | 45.7 | 45.7 | 47.5 | 47.5
| 27 | 2007-03-01 | 3 | 3 | ok | 21.9667 | 21.9667 | 21.7667 | 21.7667
| 28 |  |  |  |  | FALSE |  | FALSE
| 29 |  |  |  |  | FALSE |  | FALSE
| 30 |  |  |  |  | FALSE |  | FALSE
| 31 |  |  |  |  | FALSE |  | FALSE
| 32 |  |  |  |  | FALSE |  | FALSE
| 33 |  |  |  |  | FALSE |  | FALSE
| 34 |  |  |  |  | FALSE |  | FALSE
| 35 |  |  |  |  | FALSE |  | FALSE
| 36 |  |  |  |  | FALSE |  | FALSE
| 37 |  |  |  |  | FALSE |  | FALSE
| 38 |  |  |  |  | FALSE |  | FALSE
| 39 |  |  |  |  | FALSE |  | FALSE
| 40 |  |  |  |  | FALSE |  | FALSE
| 41 |  |  |  |  | FALSE |  | FALSE
| 42 |  |  |  |  | FALSE |  | FALSE
| 43 |  |  |  |  | FALSE |  | FALSE
| 44 |  |  |  |  | FALSE |  | FALSE
| 45 |  |  |  |  | FALSE |  | FALSE
| 46 |  |  |  |  | FALSE |  | FALSE
Extra sample data sets | 47 |  |  |  |  | FALSE |  | FALSE
| 48 |  |  |  |  | FALSE |  | FALSE
| 49 |  |  |  |  | FALSE |  | FALSE
| 50 |  |  |  |  | FALSE |  | FALSE
| 51 |  |  |  |  | FALSE |  | FALSE
| 52 |  |  |  |  | FALSE |  | FALSE
| 53 |  |  |  |  | FALSE |  | FALSE
| 54 |  |  |  |  | FALSE |  | FALSE
| 55 |  |  |  |  | FALSE |  | FALSE
| 56 |  |  |  |  | FALSE |  | FALSE
| 57 |  |  |  |  | FALSE |  | FALSE
| 58 |  |  |  |  | FALSE |  | FALSE
| 59 |  |  |  |  | FALSE |  | FALSE
| 60 |  |  |  |  | FALSE |  | FALSE
| 61 |  |  |  |  | FALSE |  | FALSE
| 62 |  |  |  |  | FALSE |  | FALSE
| 63 |  |  |  |  | FALSE |  | FALSE
| 64 |  |  |  |  | FALSE |  | FALSE
| 65 |  |  |  |  | FALSE |  | FALSE
| 66 |  |  |  |  | FALSE |  | FALSE
| 67 |  |  |  |  | FALSE |  | FALSE
| 68 |  |  |  |  | FALSE |  | FALSE
| 69 |  |  |  |  | FALSE |  | FALSE
| 70 |  |  |  |  | FALSE |  | FALSE
|  | FRM mean concentration (R-bar): | [6 empty] | 28.2885 | Regression statistics |  | Slope1 | Intercept2 | Correlation (r)
|  | Candidate mean concentration (C-bar): | [6 empty] | 28.5077 | Statistics for this test site: |  | 1.0092 | -0.042378 | 0.99651
|  | Candidate/FRM ratio: | [6 empty] | 100.8% | Limits for | Upper: | 1.1 | 1.7279
|  | Number of valid data sets: | [6 empty] | 26 | PM2.5 Class III | Lower: | 0.9 | -2 | 0.94984
|  | CCV (FRM): | [6 empty] | 0.4992 |  | Test Results (Pass/Fail): | PASS | PASS | PASS
1Multiplicative bias 2Additive bias
[10 empty] | Intercept limits | Upper | Lower
[10 empty] | PM2.5 Class II | 1.3708 | -1.5
[10 empty] | PM2.5 Class III | 1.7279 | -2
[10 empty] | PM10-2.5 Class II | 3.5 | -3.5
[10 empty] | PM10-2.5 Class III | 7 | -7
Note on G7: Validity requires that both the corresponding FRM and candidate data sets are valid [§53.35(c)(2)].
Note on K7: Means of the FRM and Candidate method measurements for valid data sets [§53.35(2) and (4), Equations 11 and 12, respectively].
Note on K80: Mean of the FRM method measurements for all valid data sets at the site, C-bar [§53.35(g)(1), Equation 17].
Note on K81: Mean of the Candidate method measurements for all valid data sets at the site, R-bar [§53.35(g)(2), Equations 18].
Note on K84: Concentration coefficient of variation (CCV) is the relative standard deviation (CV) of the valid FRM measurements for the site [§53.35(h)(2), Equation 22].
Note on P81: Calculated linear regression slope of the Candidate method versus the FRM method valid measurement pairs [§53.35(g)(3), Equation 19].
Note on Q81: Calculated linear regression intercept of the Candidate method versus the FRM method valid measurement pairs [§53.35(g)(3), Equation 20].
Note on R81: Calculated (Pearson) correlation coefficient of the Candidate method versus the FRM method valid measurement pairs [§53.35(h)(1), Equation 21].
Note on P82: [Limit specified by Table C-4.]
Note on Q82: [Limit specified by Table C-4.]
Note on P83: [Limit specified by Table C-4.]
Note on Q83: [Limit specified by Table C-4.]
Note on R83: [Limit specified by Table C-4.]
Note on O84: Test results are based only on the calculated values, not considering whether the number of valid data sets is sufficient.
Note on P84: Test result is based only on the calculated value, not considering whether the number of valid data sets is sufficient.
Note on Q84: Test result is based only on the calculated value, not considering whether the number of valid data sets is sufficient.
Note on R84: Test result is based only on the calculated value, not considering whether the number of valid data sets is sufficient.

## Summary
Summary
Applicant: |  | Example
Candidate method: |  | ACME Model XYZ123 Air monitor - PM2.5 Class III
Test site: |  | St. Louis test site - (Site location C )
Data sets | [7 empty] | Number
Valid data sets available: | [7 empty] | 26
Number of valid data sets required for Class III, location C: | [7 empty] | 23
Number of valid data sets for this test is: | [6 empty] | OK
Additional data sets needed: | [7 empty] | --
|  |  |  | Data set mean, μg/m3 |  | Data set precision, μg/m3 |  | Relative precision (CV)
Precision |  |  |  | FRM | Candidate | FRM | Candidate | FRM | Candidate
|  |  | Mean: | 28.2885 | 28.5077 | 0.78075 | 1.6563 | 3.689% | 8.837%
|  |  | Maximum: | 48.5 | 48.8333 | 1.193 | 2.616 | 8.313% | 38.97%
|  |  | Minimum: | 7.15 | 5.8 | 0.21213 | 0.45826 | 0.9926% | 1.801%
|  |  | Candidate / FRM Ratio: |  | 100.8% |  | 212.1% |  | 239.5%
|  |  |  | RMS Relative Precision for this site: |  |  |  | 4.357% | 12.23%
|  |  |  | Test requirements - PM2.5 Class III: |  |  |  | 10% | 15%
|  |  |  | Precision Test Results for site: |  |  |  | OK | PASS
Regression statistics |  |  |  |  | Slope1 | Intercept2 | Correlation (r)
Statistics for this test site: |  |  |  |  | 1.0092 | -0.042378 | 0.99651
Limits for |  |  |  | Upper: | 1.1 | 1.7279
PM2.5 Class III |  |  |  | Lower: | 0.9 | -2 | 0.94984
Test Results (Pass/Fail): |  |  |  |  | PASS | PASS | PASS
1Multiplicative bias 2Additive bias
*If chart does not plot correctly, go to the Regression sheet and click on the ▼ in the Validity column and then on "ok." If new data
are added, click "all" then "ok" to include the new data.
Calculations for the intercept limits graphic:
Precision limit calculation |  |  |  | Limits | PM2.5-II |  | PM2.5-III |  | PMc-II |  | PMc-III |  | Applicable
Size | PM2.5 |  |  | Point | PM2.5-II-S | PM2.5-II-I | PM2.5-III-S | PM2.5-III-I | PMc-II-S | PMc-II-I | PMc-III-S | PMc-III-I | S | I
Size | PM2.5 |  |  | 1 | 0.9 | 1.5 | 0.9 | 2 | 0.9 | 3.5 | 0.88 | 7 | 0.9 | 2
Cat (II/III) | III |  |  | 2 | 1.0007 | 1.5 | 0.98864 | 2 | 1.0702 | 3.5 | 1.0383 | 7 | 0.98864 | 2
Comb | PM2.5-III- |  |  | 3 | 1.1 | 0.005 | 1.1 | 0.53 | 1.1 | 1.4 | 1.12 | 2.0008 | 1.1 | 0.53
|  |  |  | 4 | 1.1 | -1.5 | 1.1 | -2 | 1.1 | -3.5 | 1.12 | -7 | 1.1 | -2
|  |  |  | 5 | 0.99934 | -1.5 | 0.98441 | -2 | 0.92979 | -3.5 | 0.93452 | -7 | 0.98441 | -2
|  |  |  | 6 | 0.9 | -0.004 | 0.9 | -0.538 | 0.9 | -1.4 | 0.88 | -2.4784 | 0.9 | -0.538
|  |  |  | 7 | 0.9 | 1.5 | 0.9 | 2 | 0.9 | 3.5 | 0.88 | 7 | 0.9 | 2
Note on H24: Calculated linear regression slope of the Candidate method versus the FRM method valid measurement pairs [§53.35(g)(3), Equation 19].
Note on I24: Calculated linear regression intercept of the Candidate method versus the FRM method valid measurement pairs [§53.35(g)(3), Equation 20].
Note on J24: Calculated (Pearson) correlation coefficient of the Candidate method versus the FRM method valid measurement pairs [§53.35(h)(1), Equation 21].
Note on H25: [Limit specified by Table C-4.]
Note on I25: [Limit specified by Table C-4.]
Note on H26: [Limit specified by Table C-4.]
Note on I26: [Limit specified by Table C-4.]
Note on J26: [Limit specified by Table C-4.]
Note on I9: At least 46 valid data sets required for site location A, 23 for site locations B, C, and D [§53.35(b)(2) and (c)(2), and Table C-5]
Note on K9: At least 46 valid data sets required for site location A, 23 for site locations B, C, and D [§53.35(b)(2) and (c)(2), and Table C-5]
Note on H13: Means of the FRM and Candidate method measurements for valid data sets [§53.35(d)(2) and (4), Equations 11 and 12, respectively].
Note on J13: Precisions as the standard deviation of the FRM and candidate method measurements, respectively. These precision values are not used directly for test purposes.
Note on L13: Precisions of the FRM and Candidate method valid data sets, RPj and CPj [§53.35(e)(1) and (f)(1), Equations13 and 15]. CV = std. dev.(Col. J or K)/mean(Col. H or I).
Note on J19: FRM precision for the test site is the root mean square (RMS) of the precisions of the FRM or Candidate data sets [§53.35(e)(2) or (f)(2), Equations 14 or 16].
Note on K19: FRM precision for the test site is the root mean square (RMS) of the precisions of the FRM data sets [§53.35(e)(2), Equation 14].
Note on L19: Candidate method precision for the test site is the root mean square (RMS) of the precisions of the candidate data sets [§53.35(f)(2), Equation 16].
Note on G20: Requirement for method precision [Table C-4].
Note on K20: Requirement for FRM method precision [Table C-4].
Note on L20: Requirement for Candidate method precision [Table C-4].
Note on K21: Test result is based only on the calculated precision, not considering whether the number of valid data sets is sufficient.
Note on L21: Test result is based only on the calculated precision, not considering whether the number of valid data sets is sufficient.
Note on J21: Test results are based only on the calculated precision, not considering whether the number of valid data sets is sufficient.
Note on C27: Test results are based only on the calculated values, not considering whether the number of valid data sets is sufficient.
Note on H27: Test result is based only on the calculated value, not considering whether the number of valid data sets is sufficient.
Note on I27: Test result is based only on the calculated value, not considering whether the number of valid data sets is sufficient.
Note on J27: Test result is based only on the calculated value, not considering whether the number of valid data sets is sufficient.