Appendix H

United States Environmental Protection Agency

National Ambient Air Monitoring Technical System Audit Form

Page intentionally left blank

Table of Contents

- General / Quality Management

  - Program Organization

  - Facilities

  - Independent Quality Assurance and Quality Control

  - Planning Documents (including QMP, QAPPs, & SOPs)

  - General Documentation Policies

  - Training

  - Corrective Action

  - Quality Improvement

  - External Performance Audits

- Network Management / Field Operations

  - Network Design

  - Changes to the Network since the last audit

  - Proposed changes to the Network

  - Field Support

    - SOPs

    - Instrument Acceptance

    - Calibration

    - Repair

    - Site and Monitor Information Form

- Laboratory Operations

  - Routine Operations

  - Quality Control

  - Laboratory Preventive Maintenance

  - Laboratory Record Keeping

  - Laboratory Data Acquisition and Handling

  - Specific Pollutants: PM10,PM 2.5 and Lead

4)   Data and Data Management

- Data Handling

b.)   Software Documentation

- Data Validation and Correction

- Data Processing

- Internal Reporting

- External Reporting

1) General / Quality Management

State/ Local / Tribal Agency Audited:

Address:

City, State, and Zip Code:

Date of Technical System Audit:

Auditor / Agency:

a) Program Organization

1) Key Individuals

1.1) Agency Director:

1.2) Ambient Air Monitoring (AAM) Network Manager:

1.3) Quality Assurance Manager:

1.4) QA Auditors:

1.5) Field Operations Supervisor / Lead:

1.6) Laboratory Supervisor:

1.7) QA Laboratory Manager:

1.8) Data Management Supervisor / Lead:

Attach an Organizational Chart :

Flow Chart:

| Key position staffing. Number of personnel available to each of the following program areas: |
| --- |
| Program Area | Number of People Primary | Number of People Backup | Vacancies | Program Area | Number of People Primary | Number of People Backup | Vacancies |
| Network Design and Siting |  |  |  | Data and Data Management |  |  |  |
| QC activities |  |  |  | Equipment repair and maintenance |  |  |  |
| QA activities |  |  |  | Financial Management |  |  |  |

| List available personnel by name and percentage of time spent on each task category |
| --- |
| Name | Network Design and siting | QC activities | QA activities | Equipment repair and maintenance | Data and Data Management | Financial Management |

Comment on the need for additional personnel, if applicable

| List personnel who have authority or are responsible for: |
| --- |
| Activity | Name | Title |
| QA Training Field/Lab |  |  |
| Grant Management |  |  |
| Purchases greater than $500 |  |  |
| Equipment and Service Contract Management |  |  |
| Staff appointment |  |  |

b) Facilities

| Identify the principal facilities where the agency conducts work that is related to air monitoring. Do not include monitoring stations but do include facilities where work is performed by contractors or other organizations. |
| --- |
| Facility AAM Function | Offices responsible for ensuring adequacy | Location | Adequate Y/N To be completed by auditor |
| Instrument repair |  |  |  |
| Certification of Standards e.g. gases, flow transfers, MFC |  |  |  |
| PM filter weighing |  |  |  |
| Data verification and processing |  |  |  |
| General office space |  |  |  |
| Storage space, short and long term |  |  |  |
| Air Toxics (Carbonyls, VOC s, Metals): |  |  |  |
| Indicate any facilities that should be upgraded. Identify by function and any suggested improvements or recommendations |
| Are facilities adequate concerning safety? Yes / No Please explain if answer is no any suggested improvements or recommendations |

| Are there any significant changes which are likely to be implemented to agency facilities within the next one to two years? Comment on agency’s needs for additional physical space (laboratory, office, storage, etc.). |
| --- |
| Facility | Function | Proposed Change - Date |

c) Independent Quality Assurance and Quality Control

1. Status of Quality Assurance Program

| Question | Yes | No | Comment |
| --- | --- | --- | --- |
| Does the agency perform QA activities with internal personnel? If no go to Section d. |  |  |  |
| Does the agency maintain a separate laboratory to support quality assurance activities? |  |  |  |
| Has the agency documented and implemented specific audit procedures separate from monitoring procedures? |  |  |  |
| Are there two levels of management separation between QA and QC operations? Please explain: |  |  |  |
| Does the agency have identifiable auditing equipment and standards (specifically intended for sole use) for audits? |  |  |  |

2. Internal Performance Audits

| Question | Yes | No | Comment |
| --- | --- | --- | --- |
| Does the agency have separate facilities to support audits and calibrations? |  |  |  |
| If the agency has in place contracts or similar agreements either with another agency or contractor to perform audits or calibrations, please name the organization and briefly describe the type of agreement. |
| If the agency does not have a performance audit SOP (included as an attachment), please describe performance audit procedure for each type of pollutant. |
| Does the agency maintain independence of audit standards and personnel? |  |  |  |
| Please provide information on certification of audit standards currently being used. Include information on vendor and internal or external certification of standards. |
| Does the agency have a certified source of zero air for performance audits? |  |  |  |
| Does the agency have procedures for auditing and/or validating performance of Meteorological monitoring? |  |  |  |
| Please provide a list of the agency’s audit equipment and age of audit equipment. |
| Is audit equipment ever used to support routine calibration and QC checks required for monitoring network operations? |  |  |  |
| If yes, please describe. |
| Are standard operating procedures (SOPs) for air monitoring available to all field personnel? |  |  |  |
| Has the agency established and has it documented criteria to define agency-acceptable audit results? |  |  |  |

| Please complete the table below with the pollutant, monitor and acceptance criteria. |
| --- |
| Pollutant | How is performance tracked (e.g., control charts) | Audit Result Acceptance Criteria |
| CO |  |  |
| O3 |  |  |
| NO2 |  |  |
| SO2 |  |  |
| PM10 |  |  |
| PM2.5 |  |  |
| Pb |  |  |
| VOCs |  |  |
| Carbonlys |  |  |
| PM2.5 speciation |  |  |
| PM10-2.5 speciation |  |  |
| PM10-2.5 FRM Mass |  |  |
| Continuous PM2.5 |  |  |
| Trace Levels (CO) |  |  |
| Trace Levels (SO2) |  |  |
| Trace Levels (NO) |  |  |
| Trace Levels (NOy) |  |  |
| Surface Meteorology |  |  |
| Others |  |  |

| Question | Yes | No | Comment |
| --- | --- | --- | --- |
| Were these audit criteria based on, or derived from, the guidance found in Volume II of the QA Handbook for Air Pollution Measurement System, Section 2.0.12? |  |  | If no, please explain. If yes, please explain any changes or assumptions made in the derivation. |
| What corrective action may be taken if criteria are exceeded? If possible, indicate two examples of corrective actions, taken within the period since the previous systems audit which are based directly on the criteria discussed above. Corrective Action # 1 Corrective Action #2 |

d) Planning Documents including QMP, QAPP, &SOP

| QMP questions | Yes | No |  |
| --- | --- | --- | --- |
| Does the agency have an EPA-approved quality management plan? |  |  |  |
| If yes, have changes to the plan been approved by the EPA? |  |  |  |
| Has the QMP been approved by EPA within the last five years? |  |  |  |
| Please provide: Date of Original Approval Date of Last Revision: Date of Latest Approval: |
| QAPP questions | Yes | No | Comment |
| Does the agency have an EPA-approved quality assurance project plan? |  |  |  |
| If yes, have changes to the plan been approved by the EPA? |  |  |  |
| Has the QAPP been reviewed by EPA annually? |  |  |  |
| Please provide: Date of Original Approval Date of Last Revision Date of Latest Approval |
| Does the agency have any revisions to your QA project plan still pending? |  |  |  |
| How does the agency verify the QA project plan is fully implemented? |  |
| How are the updates distributed? |  |
| What personnel regularly receive updates? |  |
| SOP questions |  |  |  |
| Has the agency prepared and implemented standard operating procedures (SOPs) for all facets of agency operation? |  |  |  |
| Do the SOPs adequately address ANSI/ASQC E-4 quality system required by 40 CFR 58, Appendix A? |  |  |  |
| Are copies of the SOP or pertinent sections available to agency personnel? |  |  |  |
| How does the agency verify that the SOPs are implemented as provided? |  |
| How are the updates distributed? |  |

| e) General Documentation Policies |
| --- |
| Question | Yes | No | Comment |
| Does the agency have a documented records management plan? |  |  |  |
| Does the agency have a list of files considered official records and their media type I.E. paper, electronic? |  |  |  |
| Does the agency have a schedule for retention and disposition of records? |  |  |  |
| Are records for at least three years? |  |  |  |
| Who is responsible for the storage and retrieval of records? |  |
| What security measures are utilized to protect records? |  |
| Where/when does the agency rely on electronic files as primary records? |  |
| What is the system for the storage, retrieval and backup of these files? |  |

| f) Training |
| --- |
| Question | Yes | No | Comment |
| Does the agency have a training program and training plan? |  |  |  |
| Where is it documented? |  |
| Does it make use of seminars, courses, EPA sponsored college level courses? |  |  |  |
| Are personnel cross-trained for other ambient air monitoring duties? |  |  |  |
| Are training funds specifically designated in the annual budget? |  |  |  |
| Does the training plan include: | Yes | No | Comment |
| Training requirements by position |  |  |  |
| Frequency of training |  |  |  |
| Training for contract personnel |  |  |  |
| A list of core QA related courses |  |  |  |

| Indicate below the three most recent training events and identify the personnel participating in them. |
| --- |
| Event | Dates | Participant(s) |

| Oversight of Contractors and Suppliers |
| --- |
| Question Contractors | Yes | No | Comment |
| Who is responsible for oversight of contract personnel? |  |
| What steps are taken to ensure contract personnel meet training and experience criteria? |  |
| How often are contracts reviewed and /or renewed? |  |
| Question Suppliers |  |  |  |
| Have criteria and specification been established for consumable supplies and for equipment? |  |  |  |
| What supplies and equipment have established specifications? |  |
| Is equipment from suppliers open for bid? |  |  |  |

g) Corrective Action

| Question | Yes | No | Comment |
| --- | --- | --- | --- |
| Does the agency have a comprehensive corrective action program in place and operational? |  |  |  |
| Have the procedures been documented? |  |  |  |
| As a part of the QA project plan? |  |  |  |
| As a separate standard operating procedure? |  |  |  |
| Does the agency have established and documented corrective limits for QA and QC activities? |  |  |  |
| Are procedures implemented for corrective actions based on results of the following which fall outside the established limits: |  |  |  |
| Performance evaluations |  |  |  |
| Precision goals |  |  |  |
| Bias goals |  |  |  |
| NPAP audits |  |  |  |
| PEP audits |  |  |  |
| Validation of one point QC check goals |  |  |  |
| Completeness goals |  |  |  |
| Data audits |  |  |  |
| Calibrations and zero span checks |  |  |  |
| Technical Systems Audit |  |  |  |
| Have the procedures been documented? |  |  |  |
| How is responsibility for implementing corrective actions assigned? Briefly discuss. |
| How does the agency follow up on implemented corrective actions? |
| Briefly describe recent examples of the ways in which the above corrective action system was employed to remove problems. |

| h) Quality Improvement |  |  |  |
| --- | --- | --- | --- |
| Question | Yes | No | Comment |
| What actions were taken to improve the quality system since the last TSA? |  |
| Since the last TSA do your control charts indicate that the overall data quality for each pollutant steady or improving? |  |  |  |
| For areas where data quality appears to be declining has a cause been determined? |  |  |  |
| Have all deficiencies indicted on the previous TSA been corrected? |  |  |  |
| If not explain. |
| Are there pending plans for quality improvement such as purchase of new or improved equipment, standards, or instruments? |  |  |  |

i) External Performance Audits

| Question | Yes | No | Comment |
| --- | --- | --- | --- |
| Does your agency participate in NPAP, PM2.5 PEP, and Other performance audits performed by an external party and/or using external standards. |  |  |  |
| If the agency does not participate, please explain why not: |
| Are NPAP audits performed by QA staff, site operators, calibration staff, and/or another group? |  |  |  |

National Performance Audit Program (NPAP) and Additional Audits

Does the agency participate in the National Performance Audit Program (NPAP) as required under 40 CFR 58, Appendix A? If so, identify the individual with primary responsibility for the required participation in the National Performance Audit Program.

Name:                                   Program function:

| Please complete the table below: |
| --- |
| Parameter Audited | Date of Last NPAP Audit |
| CO |  |
| O3 |  |
| SO2 |  |
| NO2 |  |
| PM10 |  |
| PM2.5 |  |
| Pb |  |
| VOCs |  |
| Carbonlys |  |

2)  Network Management/Field Operations

State/Local/Tribal Agency Audited:

Address:

City, State, and Zip Code:

Auditor / Agency:

Key Individuals

Ambient Air Monitoring Network Manager:

Quality Assurance Manager:

Field Operations Supervisor/Lead:

Field Operations Staff involved in the TSA:

| a) Network Design |
| --- |
| Complete the table below for each of the pollutants monitored as part of your air monitoring network. (Record applicable count by category.) Also indicate seasonal monitoring with an S for a Parameter/Category as appropriate. Provide the most recent annual monitoring network plan, including date of approval and AQS quicklook or if not available, network description and other similar summary of site data, including SLAMS, Other and Toxics |
| Category* | SO2 | NO2 | CO | O3 | PM10 | PM2.5 | Pb | Other (type) | Other (type) |
| NCore |  |  |  |  |  |  |  |  |  |
| SLAMS |  |  |  |  |  |  |  |  |  |
| SPM |  |  |  |  |  |  |  |  |  |
| PAMS |  |  |  |  |  |  |  |  |  |
| Total |  |  |  |  |  |  |  |  |  |

*NCore - National Core monitoring stations; SLAMS - state and local air monitoring stations; SPM - special purpose monitors; PAMS - photochemical assessment monitoring stations

| Question | Yes | No | Comment |
| --- | --- | --- | --- |
| What is the date of the most current Monitoring Network Plan? |  |
| I. Is it available for public inspection |  |  |  |
| II. Does it include the information required for each site? |  |  |  |
| AQS Site ID # |  |  |  |
| Street address and geographic coordinates |  |  |  |
| Sampling and Analysis Method(s) |  |  |  |
| Operating Schedule |  |  |  |
| Monitoring Objective and Scale of Representativeness |  |  |  |
| Site suitable/not suitable for comparison to annual PM2.5 NAAQS? |  |  |  |
| MSA, CBSA or CSA indicated as required? |  |  |  |

| Indicate by Site ID # any non-conformance with the requirements of 40 CFR 58, Appendices D and E, along with any waivers granted by the Regional Office (provide waiver documentation) |
| --- |
| Monitor | Site ID | Reason for Non-Conformance |
| SO2 |  |  |
| O3 |  |  |
| CO |  |  |
| NO2 |  |  |
| PM10 |  |  |
| PM2.5 |  |  |
| Pb |  |  |

| Question | Yes | No | Comment |
| --- | --- | --- | --- |
| Are hard copy site information files retained by the agency for all air monitoring stations within the network? |  |  |  |
| Does each station have the required information including: |  |  |  |
| AQS Site ID Number? |  |  |  |
| Photographs/slides to the four cardinal compass points? |  |  |  |
| Startup and shutdown dates? |  |  |  |
| Documentation of instrumentation? |  |  |  |
| Who has custody of the current network documents |  |  | Name: Title: |
| Does the current level of monitoring effort, station placement, instrumentation, etc., meet requirements imposed by current grant conditions? |  |  |  |
| How often is the network siting reviewed? |  |  | Frequency: Date of last review: |
| Are there any issues |  |  |  |
| Do any sites vary from the required frequency in 40 CFR 58.12? |  |  |  |
| Does the number of collocated monitoring stations meet the requirements of 40 CFR 58 Appendix A? |  |  |  |

| Changes to the Network since the last audit |
| --- |
| W What is the date of the most recent network assessment? (Provide copy) Are all SLAMS parameters included? Any Others? |
| Please provide information on any site changes since the last audit |
| Pollutant | Site ID | Site Address | Site Added/Deleted/Relocated | Reason (Assessment, lost lease, etc. Provide documentation of reason for each site change.) |

| c) Proposed changes to the Network |
| --- |
| W Are future network changes proposed? |
| Please provide information on proposed site changes, including documentation of the need for the change and any required approvals |
| Pollutant | Site ID | Site Address | Site to be Added/Deleted/Relocated | Reason (Assessment, lost lease, etc. Provide documentation of reason for each site change.) |

| d) Field Support |
| --- |
| Question | Yes | No | Comment |
| On average, how often are most of your stations visited by a field operator? |  |  | ______ per _______ |
| Is this visit frequency consistent for all reporting organizations within your agency? |  |  |  |
| On average, how many stations does a single operator have responsibility for? |  |
| How many of the stations of your SLAMS/NCORE network are equipped with sampling manifolds? |  |
| Do the sample inlets and manifolds meet the requirements for through the probe audits? |  |
| I. Briefly describe most common manifold type |  |
| II. Are Manifolds cleaned periodically |  |  | How often? |
| III. If the manifold is cleaned, what is used to perform cleaning |  |
| IV. Are manifold(s) equipped with a blower |  |  |  |
| V. Is there sufficient air flow through the manifold at all times? |  |  | Approximate air flow: |
| VI. How is the air flow through the manifold monitored? |  |
| VII. Is there a conditioning period for the manifold after cleaning? |  |  | Length of time: |
| VIII. What is the residence time? |  |
| Sampling lines: 1) What material is used for instrument sampling lines? 2) How often are lines changed? |  |
| Do you utilize uninterruptable power supplies or backup power sources at your sites? |  |  |  |
| What instruments or devices are protected? |  |  |  |

i). SOPs

| Question | Yes | No | Comment |
| --- | --- | --- | --- |
| Is the documentation of monitoring SOPs complete? |  |  |  |
| Are any new monitoring SOPs needed? |  |  |  |
| Are such procedures available to all field operations personnel? |  |  |  |
| Are SOPs that detail operations during episode monitoring prepared and available to field personnel? |  |  |  |
| Are SOPs based on the framework contained in Guidance for Preparing Standard Operating Procedures EPA QA/G-6? |  |  |  |

Please complete the following table:

| Pollutant Monitored | Date of Last SOP Review | Date of Last SOP Revision |
| --- | --- | --- |
| SO2 |  |  |
| NO2 |  |  |
| CO |  |  |
| O3 |  |  |
| PM10 |  |  |
| PM2.5 FRM mass |  |  |
| Pb |  |  |
| PM2.5 speciation |  |  |
| PM10-2.5 FRM mass |  |  |
| PM10-2.5speciation |  |  |
| Continuous PM2.5 mass |  |  |
| Trace levels (CO) |  |  |
| Trace levels (SO2) |  |  |
| Trace levels (NO) |  |  |
| Trace levels (NOy) Total reactive nitrogen |  |  |
| Surface Meteorology Wind speed and direction, temperature, RH, precipitation and solar radiation |  |  |
| Others |  |  |

ii). Instrument Acceptance

Has your agency obtained necessary waiver provisions to operate equipment which does not meet the effective reference and equivalency requirements?      List all waivers.

Please list instruments in your inventory

| Pollutant | Number | Make and Models | Reference or Equivalent number |
| --- | --- | --- | --- |
| SO2 |  |  |  |
| NO2 |  |  |  |
| CO |  |  |  |
| O3 |  |  |  |
| PM10 |  |  |  |
| PM2.5 |  |  |  |
| Pb |  |  |  |
| Multi gas calibrator |  |  |  |
| PM2.5 speciation |  |  |  |
| PM10-2.5 speciation |  |  |  |
| PM10-2.5 FRM mass |  |  |  |
| Continuous PM2.5 mass |  |  |  |
| Trace levels (CO) |  |  |  |
| Trace levels (SO2) |  |  |  |
| Trace levels (NO) |  |  |  |
| Trace levels (NOy) |  |  |  |
| Surface Meteorology |  |  |  |
| Others |  |  |  |

Please comment briefly and prioritize your currently identified instrument needs.

| Question | Yes | No | Comment |
| --- | --- | --- | --- |
| Are criteria established for field QC equipment? |  |  |  |
| Are criteria established for field QC gas standards? |  |  |  |

iii) Calibration

| Please indicate the frequency of multi point calibrations. |
| --- |
| Pollutant | Frequency | Name of Calibration Method |

| Question | Yes | No | Comment |
| --- | --- | --- | --- |
| Are field calibration procedures included in the document? SOPs? |  |  | Location (site, lab etc.): |
| Are calibrations performed in keeping with the guidance in section Vol II of the QA Handbook for Air Pollution Measurement Systems? |  |  | If no, why not? |
| Are calibration procedures consistent with the operational requirements of Appendices to 40 CFR 50 or to analyzer operation/instruction manuals? |  |  | If no, why not? |
| Have changes been made to calibration methods based on manufacturer’s suggestions for a particular instrument? |  |  |  |
| Do standard materials used for calibrations meet the requirements of appendices to 40 CFR 50 (EPA reference methods) and Appendix A to 40 CFR 58 (traceability of materials to NIST-SRMs or CRMs)? |  |  | Comment on deviations |
| Are all flow-measurement devices checked and certified? |  |  |  |
| Additional comments: |
| Please list the authoritative standards used for each type of flow measurement, indicate the certification frequency of standards to maintain field material/device credibility. |
| Flow Device | Primary Standard | Frequency of Certification |
| HiVol orifice |  |  |
| Streamline |  |  |
| TriCal |  |  |
| BIOS |  |  |
| DeltaCal |  |  |
| Gilibrators |  |  |
| Where do field operations personnel obtain gaseous standards? |  |
| Are those standards certified by: |  |  |  |
| The agency laboratory? |  |  |  |
| EPA/NERL standards laboratory? |  |  |  |
| A laboratory separate from this agency’s but part of the same reporting organization? |  |  |  |
| The vendor? |  |  |  |
| Other (describe). |  |  |  |
| How are the gas standards verified after receipt? |  |
| How are flow measurement devices certified? |  |  |
| Please provide copies of certifications of all standards currently in use from your master and/or satellite standard certification logbooks (i.e., chemical standards, ozone standards, flow standards, and zero air standards) . |  |  |  |
| What equipment is used to perform calibrations (e.g., dilution devices) and how is the performance of this equipment verified? |  |
| Does the documentation include expiration date of certification? |  |  |  |
| Reference to primary standard used? |  |  |  |
| What traceability is used? |  |
| Please attach an example of recent documentation of traceability |  |  |  |
| Is calibration equipment maintained at each station? |  |  |  |
| How is the functional integrity of this equipment documented? |  |
| Who has responsibility for maintaining field calibration standards? |
| Please list the authoritative standards and frequency of each type of dilution, permeation and ozone calibrator and indicate the certification frequency... |
| Calibrator | Primary Standard | Frequency of CalibrationCertification |
| Permeation calibrator flow controller |  |  |
| Permeation calibrator temperature |  |  |
| Dilution calibrator air and gas flow controllers |  |  |
| Field/working standard photometer |  |  |
| Ozone generator |  |  |

| o) Please identify station standards for gaseous pollutants at representative air monitoring stations (attach additional sheets as appropriate): |
| --- |
| Parameter | Station(s) | Identification of Standard(s) | Recertification Date(s) |
| CO |  |  |  |
| NO2 |  |  |  |
| SO2 |  |  |  |
| O3 |  |  |  |

iv) Repair

a)  Who is responsible for performing preventive maintenance?

b)  Is special training provided them for performing preventive maintenance?   Briefly comment on background or courses.

c) Is this training routinely reinforced? Yes        No

If no, why not?

d)  What is your preventive maintenance schedule for each type of field instrumentation?

e) If preventive maintenance is MINOR, it is performed at (check one or more): field station       , headquarters facilities       , equipment is sent to manufacturer

f) If preventive maintenance is MAJOR, it is performed at (check one or more): field station      , headquarters facilities       , equipment is sent to manufacturer

g) Does the agency have service contracts or agreements in place with instrument manufacturers?  Indicate below or attach additional pages to show which instrumentation is covered?

h) Comment briefly on the adequacy and availability of the supply of spare parts, tools and manuals available to the field operator to perform any necessary maintenance activities. Do you feel that this is adequate to prevent any significant data loss?

i) Is the agency currently experiencing any recurring problem with equipment or manufacturer(s)?  If so, please identify the equipment or 	manufacturer, and comment on steps taken to remedy the problem.

j) Have you lost any data due to repairs in the last 2 years?

More than 24 hours?

More than 48 hours?

More than a week?

k)  Explain any situations where instrument down time was due to lack of preventive maintenance of unavailability of parts.

RECORD KEEPING

| Question | Yes | No | Comment |
| --- | --- | --- | --- |
| What type of station logbooks are maintained at each monitoring station? (maintenance logs, calibration logs, personal logs, etc.) |  |
| What information is included in the station logbooks? |  |
| Who reviews and verifies the logbooks for adequacy of station performance? |  |
| How is control of logbook maintained? |  |
| Where is the completed logbook archived? |  |
| What other records are used? |  |
| Zero span record? |  |  |  |
| Gas usage log? |  |  |  |
| Maintenance log? |  |  |  |
| Log of precision checks? |  |  |  |
| Control charts? |  |  |  |
| A record of audits? |  |  |  |
| Please describe the use and storage of these documents. |
| Are calibration records or at least calibration constants available to field operators? |  |  |  |
| Please attach an example field calibration record sheet to this questionnaire. |

##### V) Site/Monitor Information Form

PQAO_______________________________________________________

AQS Site Name

AQS Site Number

Agency Site Name/No.

(if different than AQS Site Name/Number)

Site Address

City & County

Site Coordinates

(specify lat/long or UTM)

## Site Elevation (m)

Criteria Pollutants Monitored

Other Parameters

Nearst Meterological Site

(‘on site’ is met tower present at this site)

Photographs to and from each cardinal direction attached?

(Yes or No)

Name(s) of Report Preparer(s)

Name(s) of Auditors ____________________________________________

Date

## Phone Number

Site Map

Draw map of site and surrounding terrain and features, up to 100 meters.

Map notes

Monitor Information

### Pollutants

| Manufacturer |  |  |  |
| --- | --- | --- | --- |
| Model |  |  |  |
| Serial number |  |  |  |
| Scale of representation MICro, MIDdle, Neighborhood, Urban |  |  |  |
| Objective (Population, Max concentration, Background, Transport) |  |  |  |
| Height of probe above ground(m) |  |  |  |
| Distance from obstruction (m) |  |  |  |
| Type of obstruction (Wall, Tree, etc) |  |  |  |
| Distance from roadway (m) |  |  |  |
| Unrestricted airflow (Yes, No) |  |  |  |
| Designation (NCore, SLAMS,etc) |  |  |  |
| Siting Criteria Met (Yes, No) |  |  |  |

| Manufacturer |  |  |  |
| --- | --- | --- | --- |
| Model |  |  |  |
| Serial number |  |  |  |
| Scale of representation MICro, MIDdle, Neighborhood, Urban |  |  |  |
| Averaging time 1-, 8-, 24-hour |  |  |  |
| Objective (Population, Max concentration, Background, Transport) |  |  |  |
| Height of probe above ground(m) |  |  |  |
| Distance from obstruction (m) |  |  |  |
| Type of obstruction (Wall, Tree, etc) |  |  |  |
| Distance from roadway (m) |  |  |  |
| Unrestricted airflow (Yes, No) |  |  |  |
| Designation (NCore, SLAMS,etc) |  |  |  |
| Siting Criteria Met (Yes, No) |  |  |  |

Insert additional copies of table as needed

Area Information

| Street Name | Traffic Count (Vehicles/day) |
| --- | --- |

| Direction | Predominant Land Use (Industry, Residential, Commercial or Agriculture) |
| --- | --- |
| North |  |
| East |  |
| South |  |
| West |  |

| Direction | Obstructions | Height (m) | Distance (m) |
| --- | --- | --- | --- |
| North |  |  |  |
| East |  |  |  |
| South |  |  |  |
| West |  |  |  |

Note: This table is for large obstructions that affect the entire site, such as large clusters of trees or entire buildings.  Individual obstructions, such as walls, single trees, other monitors, etc, should be entered in the Monitor Information table.

| Direction | Topographic Features (hills, valleys, rivers, etc.) | General Terrain (flat, rolling, rough) |
| --- | --- | --- |
| North |  |  |
| East |  |  |
| South |  |  |
| West |  |  |

Comments

3) LABORATORY OPERATIONS

State/Local/Tribal Agency Audited:

City, State, and Zip Code:

Date of Technical System Audit:

Auditor / Agency:

Key Individuals

Laboratory Manager:

Laboratory Supervisor:

Quality Assurance Manager:

Laboratory Staff involved in the TSA:

a) Routine Operations

What analytical methods are employed in support of your air monitoring network?

Analysis		Name or Description of Method

PM10

PM2.5

Pb

Others (list by pollutant)

Please describe areas where there have been difficulties meeting the regulatory requirements for any of the above analytical methods.

| In the table below, please identify the current versions of written methods, supplements, and guidelines that are used in your agency. |
| --- |
| Analysis | Documentation of Method |
| PM10 |  |
| PM2.5 |  |
| Pb |  |
| Others (list by pollutant) |  |

| Question | Yes | No | Comment |
| --- | --- | --- | --- |
| Were procedures for the methods listed above included in the agency(s QA Project Plan or SOP(s and were reviewed by EPA? Also, are SOP’s easily/readily accessible for use and reference? |  |  |  |
| Does you lab have sufficient instrumentation to conduct analyses? |  |  |  |

d) Please describe needs for laboratory instrumentation.

b) Laboratory Quality Control

Please identify laboratory standards used in support of the air monitoring program, including standards which may be kept in an analytical laboratory and standards which may be kept in a field support area or quality assurance laboratory that is dedicated to the air monitoring program (attach additional sheets if appropriate):

| Parameter | Location of Standards | Laboratory Standard | Recertification Date | Primary Standard* |
| --- | --- | --- | --- | --- |
| CO |  |  |  |  |
| NO2 |  |  |  |  |
| SO2 |  |  |  |  |
| O3 |  |  |  |  |
| Weights |  |  |  |  |
| Temperature |  |  |  |  |
| Moisture |  |  |  |  |
| Barometric Pressure |  |  |  |  |
| Flow |  |  |  |  |
| Other Flow Standard |  |  |  |  |
| Lead |  |  |  |  |
| Other |  |  |  |  |

*Standards to which the laboratory standards can be traced.

| Question | Yes | No | Comment |
| --- | --- | --- | --- |
| Are all chemicals and solutions clearly marked with an indication of shelf life? |  |  |  |
| Are chemicals removed and properly disposed of when shelf life expires? |  |  |  |
| Are only ACS grade chemicals used by the laboratory? |  |  |  |
| e) Comment on the traceability of chemicals used in the preparation of calibration standards |

| Question | Yes | No | Comment |
| --- | --- | --- | --- |
| Does the laboratory purchase standard solutions such as those for use with lead or other metals analysis? |  |  |  |
| Are all calibration procedures documented? |  |  |  |
| If answer “yes” to (f), please describe the following: Title of the document: Revision number: Where the document is: |
| Are at least one duplicate, one blank, and one standard or spike included with a given analytical batch? |  |  |  |
| Briefly describe the laboratory(s use of data derived from blank analyses. |
| Are criteria established to determine whether a blank data are acceptable? |  |  |  |
| How frequently and at what concentration ranges does the lab perform duplicate analysis? What constitutes an acceptable agreement? Please comment in the space below. |
| Please describe how the lab use data obtained from spiked samples, including the acceptance criteria (e.g., acceptable percent recovery). |

| Question | Yes | No | Comment |
| --- | --- | --- | --- |
| Does the laboratory routinely include samples of reference material within an analytical batch? |  |  |  |
| If yes, indicate frequency, level, and material used. |
| Are mid-range standards included in analytical batches? |  |  |  |
| Please describe the frequency, level and compound used in the space provided below. |
| Are criteria for real time quality control established that are based on the results obtained for the mid-range standards discussed above? |  |  |  |
| If yes, briefly discuss them below or indicate the document in which they can be found. |
| Are appropriate acceptance criteria for each type of analysis documented ? |  |  |  |

c) Laboratory Preventive Maintenance

| Question | Yes | No | Comment |
| --- | --- | --- | --- |
| For laboratory equipment, who has the responsibility for performing preventive maintenance? |
| Is most maintenance performed in the lab? |  |  |  |
| Is a maintenance log maintained for each major laboratory instrument? |  |  |  |
| Are service contracts in place for major analytical instruments? |  |  |  |

d) Laboratory Record Keeping

| Question | Yes | No | Comment |
| --- | --- | --- | --- |
| Are all samples that are received by the laboratory logged in? |  |  |  |
| Discuss sample routing and special needs for analysis (or attach a copy of the latest SOP which covers this). Attach a flow chart if possible. |
| Are log books kept for all analytical laboratory instruments? |  |  |  |
| Are there log books or other records that indicate the checks made on materials and instruments such as weights, humidity indicators, balances, and thermometers? |  |  |  |
| Identify type of record, acceptable/non-acceptable |
| Are log books maintained to track the preparation of filters for the field? |  |  |  |
| Are they current? |  |  |  |
| Do they indicate proper use of conditioning? |  |  |  |
| Weighings? |  |  |  |
| Stamping and numbering? |  |  |  |
| Are log books kept which track filters returning from the field for analysis? |  |  |  |
| How are data records from the laboratory archived? Where? Who has the responsibility? Person Title How long are records kept? Years |
| Does a chain-of-custody procedure exist for laboratory samples? |  | . |  |
| If yes, indicate date, title and revision number where it can be found |

e) Laboratory Data Acquisition and Handling

| Question | Yes | No | Comment |
| --- | --- | --- | --- |
| Identify those laboratory instruments which make use of computer interfaces directly to record data. Which ones use strip charts? Integrators? |
| Are QC data readily available to the analyst during a given analytical run? |  |  |  |
| What is the laboratory(s capability with regard to data recovery? In case of problems, can they recapture data or are they dependent on computer operations? Discuss briefly. |
| Has a user(s manual been prepared for the automated data acquisition instrumentation? |  |  |  |
| Please provide below a data flow diagram which establishes, by a short summary flow chart: transcriptions, validations, and reporting format changes the data goes through before being released by the laboratory. |

f) Specific Pollutants: PM10, PM2.5 and Lead

| Question | Yes | No | Comment |
| --- | --- | --- | --- |
| PM10 and PM2.5 |  |  |  |
| Does the agency use filters supplied by EPA? |  |  |  |
| Do filters meet the specifications in 40 CFR 50? |  |  |  |
| Are filters visually inspected via strong light from a view box for pinholes and other imperfections? |  |  |  |
| Where does the laboratory keep records of the serial numbers of filters? |
| Are unexposed filters equilibrated in controlled conditioning environment which meets or exceeds the requirements of 40 CFR 50? |  |  |  |
| Are the temperature and relative humidity of the conditioning environment monitored? |  |  |  |
| Are the temperature and humidity monitors calibrated? |  |  |  |
| Are balances checked with Class S or Class M weights each day when they are used? |  |  |  |
| Is the balance check information placed in QC log book? |  |  |  |
| To what sensitivity are filter weights recorded? |
| Are filter serial numbers and tare weights recorded in a bound notebook? |  |  |  |
| Are filters packaged for protection while transporting to and from the monitoring stations? |  |  |  |
| How often are filter samples collected? (Indicate the average elapsed time in hours between end of sampling and laboratory receipt.) |
| In what medium are field measurements recorded (e.g., in a log book, on a filter folder, or on standard forms)? |
| Are exposed filters reconditioned for at least 24 hrs in the same conditioning environment as for unexposed filters? |  |  |  |
| Briefly describe how exposed filters are prepared for conditioning. |
| Briefly describe how exposed filters are stored after being weighed. |
| Are blank filters reweighed? |  |  |  |
| Are chemical analyses performed on filters? |  |  |  |
| LEAD |  |  |  |
| Is analysis for lead being conducted using atomic absorption spectrometry with air acetylene flame? |  |  | If not, has the agency received an equivalency designation of their procedure? |
| Is either the hot acid or ultrasonic extraction procedure being followed precisely? |  |  | Which? |
| Is Class A borosilicate glassware used throughout the analysis? |  |  |  |
| Is all glassware cleaned with detergent, soaked and rinsed three times with distilled or deionized water? |  |  |  |
| If extracted samples are stored, are linear polyethylene bottles used? |  |  |  |
| Are all batches of glass fiber filters tested for background lead content? |  |  |  |
| At a rate of 20 to 30 random filters per batch of 500 or greater? |  |  | Indicate rate. |
| Are ACS reagent grade HNO3 and HCl used in the analysis? |  |  |  |
| Is a calibration curve available having concentrations that cover the linear absorption range of the atomic absorption instrumentation? |  |  |  |
| Is the stability of the calibration curve checked by alternately remeasuring every 10th sample a concentration ( 1(g Pb/ml; ( 10 (g Pb/ml? |  |  |  |

4) DATA AND DATA MANAGEMENT

State/Local/Tribal Agency Audited:

City, State, and Zip Code:

Date of Technical System Audit:

Auditor / Agency:

Key Individuals

Data Manager:

Data Supervisor:

Quality Assurance Manager:

a) Data Handling

| Question | Yes | No | Comment |
| --- | --- | --- | --- |
| Is there a procedure, description, or a chart which shows a complete data sequence from point of acquisition to point of submission of data to EPA? |  |  |  |
| Please provide below a data flow diagram indicating both the data flow within the reporting organization. |
| Are procedures for data handling (e.g., data reduction, review, etc.) documented? |  |  |  |
| In what media (e.g., diskette, data cartridge, or telemetry) and formats do data arrive at the data processing location? Please list below. |
| Category of Data (by Pollutant) | Data Media and Formats |
| How often are data received at the processing location from the field sites and laboratory? |
| Is there documentation accompanying the data regarding any media changes, transcriptions, or flags which have been placed into the data before data are released to agency internal data processing? |  |  | . |
| Describe the type of documentation |
| How data are actually entered to the computer system (e.g., computerized transcription (copy from disk or data transfer device), manual entry, digitization of strip charts, or other)? |
| For manual data, is a double-key entry system used? |  |  |  |

b)  Software Documentation

| Question | Yes | No | Comment |
| --- | --- | --- | --- |
| Does your agency use any AQS Manual? |  |  |  |
| Does your agency use any Air Now Manual? |  |  |  |
| If yes, list the title of manual used including the , version number and date published |
| Does the agency have information on the reporting of precision and accuracy data available? |  |  |  |
| What are the origins of the software used to prepare air monitoring data for release into the AQS and AirNow database? Please list the documentation for the software currently in use for data processing, including the names of the software packages, vendor or author, revision numbers, and the revision dates of the software. |
| What is the recovery capability in the event of a significant computer problem ( i.e., how much time and data would be lost)? |
| Has your agency tested the data processing software to ensure its performance of the intended function are consistent with the QA Handbook, Volume II, and Section 14.0? |  |  |  |
| Does your agency document software tests? |  |  | . |
| If yes, provide the documentation |

c) Data Validation and Correction

| Question | Yes | No | Comment |
| --- | --- | --- | --- |
| Have your agency established and document the validation criteria ? |  |  | If yes, indicate document where such criteria can be found (title, revision date). |
| Does documentation exist on the identification and applicability of flags (i.e., identification of suspect values) within the data as recorded with the data in the computer files? |  |  |  |
| Does your agency document the data validation criteria including limits for values such as flow rates, calibration results, or range tests for ambient measurements? |  |  |  |
| If yes, please describe what action the data validator will take if he/she find data with limits exceeded (e.g., flags, modifies, or delete, etc.) . |
| (ii) If yes, give examples to illustrate actions taken when limits were exceeded. |
| Please describe how changes made to data that were submitted to AQS and AirNow are documented. |
| Who has signature authority for approving corrections? Name Program function _____________________________ |
| What criteria are used to determine a data point be deleted? Discuss briefly. |
| What criteria are used to determine if data need to be reprocessed? Discuss. |
| Are corrected data resubmitted to the issuing group for cross-checking prior to release? |  |  |  |

d) Data Processing

| Question | Yes | No | Comment |
| --- | --- | --- | --- |
| Does the agency generate data summary reports? |  |  |  |

| Please list at least three reports routinely generated, including the information requested below. |
| --- |
| Report Title | Distribution | Period Covered |

| Question | Yes | No | Comment |
| --- | --- | --- | --- |
| How often are data submitted to AQS and AirNow? |
| Briefly comment on difficulties the agency may have encountered in coding and submitting data following the guidance of the AQS guidelines? |
| Does the agency routinely request a hard copy printout on submitted data from AQS? |  |  |  |
| Are records kept for at least 3 years by the agency in an orderly, accessible form? |  | . |  |
| If yes, does this include raw data , calculation , QC data , and reports ? If no, please comment |
| Has your agency submitted data along with the appropriate calibration equations used to the processing center? |  |  |  |
| Are concentrations of pollutants other than PM2.5 corrected to EPA standard temperature and pressure conditions (i.e.,298(K, 760 mm Hg) before input to AQS, and concentrations of PM2.5 reported to AQS under actual (volumetric) conditions? |  |  | : |
| Are audits on data reduction procedure performed on a routine basis? |  |  |  |
| If yes, at what frequency? |
| Are data precision and accuracy checked each time they are calculated, recorded, or transcribed to ensure that incorrect values are not submitted to EPA? |  |  |  |

e) Internal Reporting

| What internal reports are prepared and submitted as a result of the audits required under 40 CFR 58, Appendix A? |
| --- |
| Report Title | Frequency |

| What internal reports are prepared and submitted as a result of precision checks also required under 40 CFR 58, Appendix A? |
| --- |
| Report Title | Frequency |

| Question | Yes | No | Comment |
| --- | --- | --- | --- |
| Do either the audit or precision check reports indicated include a discussion of corrective actions initiated based on audit or precision check results? |  |  |  |

| Who has the responsibility for the calculation and preparation of data summaries? To whom are such summaries delivered? |
| --- |
| Name | Title | Type of Report | Recipient |

f) External Reporting

| For the current calendar year or portion thereof which ended at least 90 calendar days prior to the receipt of this questionnaire, please provide the following percentages for required data submitted on time. |
| --- |
| Percent Submitted on Time* Period Covered: |
| Monitoring Qtr. | SO2 | CO | O3 | NO2 | PM10 | PM2.5 | Pb |
| 1 (Jan 1 - March 31) |  |  |  |  |  |  |  |
| 2 (Apr 1 - June 30) |  |  |  |  |  |  |  |
| 3 (July 1 - Sept. 30) |  |  |  |  |  |  |  |
| 4 (Oct.1 - Dec. 31) |  |  |  |  |  |  |  |

*"On time" = within 90 calendar days after the end of the quarter in which the data were collected.

| For the same period, what fraction of the stations (by pollutant) reported less than 75% of the data (adjusted for seasonal monitoring and site start-ups and terminations)? |
| --- |
| Percent of Stations <75% Data Recovery Period Covered: |
| Monitoring Qtr. | SO2 | CO | O3 | NO2 | PM10 | PM2.5 | Pb |
| 1 (Jan 1 - March 31) |  |  |  |  |  |  |  |
| 2 (Apr 1 - June 30) |  |  |  |  |  |  |  |
| 3 (July 1 - Sept. 30) |  |  |  |  |  |  |  |
| 4 (Oct.1 - Dec. 31) |  |  |  |  |  |  |  |

Identify the individual within the agency with the responsibility for reviewing and releasing the data.

Name                                                                                                  Program function

| Question | Yes | No | Comment |
| --- | --- | --- | --- |
| Does your agency report the Pollutant Standard Index? |  |  |  |
| Has your agency submitted its annual data summary report (as required in 40 CFR 58.26)? |  |  |  |
| If yes, did your agency(s annual report include the following: |  |  |  |
| Data summary required in Appendix F? |  |  |  |
| Annual precision and accuracy information described in Section 5.2 of Appendix A? |  |  |  |
| Location, date, pollution source and duration of all episodes reaching the significant harm levels? |  |  |  |
| Is Data Certification signed by a senior officer of your agency? |  |  |  |