QA Handbook Volume II, Appendix D

Revision No. 1

Date:12/08

Page 
 of 30
QA Handbook Volume II, Appendix D

Revision No. 1

Date:12/08

Page 
 of 30

Appendix D  SEQ CHAPTER \h \r 1
Measurement Quality Objectives and Validation Templates
In June 1998, a workgroup was formed to develop a procedure that could be used by State and locals that would provide for a consistent validation of PM2.5 mass concentrations across the US.  The workgroup included personnel from the monitoring organizations, EPA Regional Offices, and OAQPS who are involved with assuring the quality of PM2.5 mass and was headed by a State and local representative.  The workgroup developed three tables of criteria where each table has a different degree of implication about the quality of the data.  The criteria included on the tables are from 40 CFR Part 50 Appendices L and N, 40 CFR Part 58 Appendix A, Method 2.12, and a few criteria that are neither in CFR nor Method 2.12.  Upon completion and use of the table, it was decided that a “validation template” should be developed for all the criteria pollutants.
One of the tables has the criteria that must be met to ensure the quality of the data.  An example criterion is that the average flow rate for the sampling period must be maintained to within 5% of 16.67 liters per minute.  The second table has the criteria that indicate that there might be a problem with the quality of the data and further investigation is warranted before making a determination about the validity of the sample or samples. An example criterion is that the field filter blanks should not change weight by more than 30 g between weighings.  The third table has criteria that indicate a potentially systematic problem with the environmental data collection activity.  Such systematic problems may impact the ability to make decisions with the data.  An example criterion is that at least 75% of the scheduled samples for each quarter should be successfully collected and validated.

To determine the appropriate table for each criterion, the members of the workgroup considered how significantly the criterion impact the resulting concentration.  This was based on experience from workgroup members, experience from non-workgroup members, and feasibility of implementing the criterion.  

Criteria that were deemed critical to maintaining the integrity of a sample or group of samples were placed on the first table.  Observations that do not meet each and every criterion on the Critical Criteria Table should be invalidated unless there are compelling reason and justification for not doing so.  Basically, the sample or group of samples for which one or more of these criteria are not met is invalid until proven otherwise.  The cause of not operating in the acceptable range for each of the violated criteria must be investigated and minimized to reduce the likelihood that additional samples will be invalidated.

Criteria that are important for maintaining and evaluating the quality of the data collection system are included on the second table, the Operational Evaluations Table.  Violation of a criterion or a number of criteria may be cause for invalidation.  The decision should consider other quality control information that may or may not indicate the data are acceptable for the parameter being controlled.  Therefore, the sample or group of samples for which one or more of these criteria are not met is suspect unless other quality control information demonstrates otherwise.  The reason for not meeting the criteria MUST be investigated, mitigated or justified.

Finally, those criteria which are important for the correct interpretation of the data but do not usually impact the validity of a sample or group of samples are included on the third table, the Systematic Issues Table.  For example, the data quality objectives are included in this table.  If the data quality objectives are not met, this does not invalidate any of the samples but it may impact the error rate associated with the attainment/non-attainment decision.

Following are the tables.  For each criterion, the tables include (1) the operational range that is acceptable, (2) the frequency with which compliance is to be evaluated, (3) the number of samples that are impacted if violation of a criterion occurs (possible values include single filters, a batch of filters, or a group of filters from a specific instrument);.(4) sections of 40 CFR and (5) Method 2.12 that describe the criterion. The table also indicates whether samples violating the criterion must be flagged before entering them into AQS.

This validation template has been developed based on the current state of knowledge.  The template should evolve as new information is discovered about the impact of the various criterion on the error in the resulting mass estimate.  Interactions of the criteria, whether synergistic or antagonistic, should also be incorporated when the impact of these interactions becomes quantified.  Due to the potential misuse of invalid data, data that are invalidated will not be uploaded to AQS but should be retained on the monitoring organizations local database.  This data will be invaluable to the evolution of the validation template.

PM10  Note of Caution

The validation templates for PM10 get complicated because PM10 is required to be reported at standard temperature and pressure (STP) for comparison to the NAAQS (and follow 40 CFR Part 50 App J) and at local conditions if using it to monitor for PM10-2.5 (and follow 40 CFR Part 50 App O) in addition PM10 is measured with filter based sampling techniques as well as with automated methods.  The validation templates developed for PM10 try to accommodate these differences but monitoring organizations are cautioned to review the operations manual for the monitors/samplers they use and augment the validation template with QC information specific to their method.

Ozone Validation Template

	Requirement
	Frequency
	Acceptance Criteria
	Information /Action
	

	CRITICAL CRITERIA-Ozone

	One Point QC Check
Single analyzer
	1/ 2 weeks
	< +7% (percent difference)
	0.01 - 0.10 ppm

Relative to routine concentrations

40 CFR Part 58 App A Sec 3.2
	

	Zero/span check 


	1/ 2  weeks
	Zero drift < + 2% of full scale

Span drift < + 7 %
	
	

	OPERATIONAL CRITERIA - Ozone

	Shelter Temperature
	
	
	
	

	   Temperature range
	Daily

(hourly  values)
	20 to 30 C.  (Hourly ave)

or

per manufacturers specifications if designated  to a wider temperature range
	Generally the 20-30  C range will apply but  the most restrictive operable range of the instruments in the shelter may also be used as guidance
	

	   Temperature Control
	Daily (hourly values)
	< + 2o C SD over 24 hours
	
	

	   Temperature Device Check
	2/year
	+ 2o C  of standard
	
	

	Precision(using 1-point QC checks)
	Calculated annually and as appropriate for design value estimates
	90% CL CV < 7%
	90% Confidence Limit of coefficient of variation. 40 CFR Part 58 App A sec 4.1.2
	

	Bias (using 1-point QC checks)
	Calculated annually and as appropriate for design value estimates
	95% CL <  + 7%
	95% Confidence Limit of absolute bias estimate. 40 CFR Part 58 App A sec 4.1.3
	

	Annual Performance Evaluation
	
	
	
	

	Single analyzer
	Every site 1/year  25 % of sites quarterly


	Percent difference of  audit levels 3-10 <  +15% 
Audit levels 1&2 + 1.5 ppb difference or  +15%
	3 consecutive audit concentration not including zero.  40 CFR Part 58 App A sec 3.2.2
	

	Primary QA Organization (PQAO)
	annually
	95% of audit percent differences fall within the one point QC check 95% probability intervals  at PQAO level of aggregation
	40 CFR Part 58 App A sec 4.1.4
	

	Federal Audits (NPAP)
	1/year at selected sites 20% of sites audited
	Mean absolute difference + 10%
	40 CFR Part 58 App A sec 2.4
	

	State audits
	1/year
	State requirements
	
	

	Verification/Calibration
	Upon receipt/adjustment/repair/ installation/moving
1/6 months if manual zero/span performed biweekly

1/year if continuous zero/span performed daily
	All points within +  2 % of full scale of best-fit straight line

Linearity error <5%
	Multi-point calibration  (0 and 4 upscale points) 40 CFR Part 50 App D sec 5.2.3
	

	Zero Air
	
	Concentrations below LDL
	
	

	Gaseous Standards
	
	NIST Traceable

(e.g., EPA Protocol Gas)
	40 CFR Part  58 App A sec 2.6.1
	

	Zero Air Check
	1/year
	Concentrations below LDL
	
	

	Ozone Local primary standard
	
	
	
	

	   Certification/recertification  to 

   Standard Reference Photometer 
	1/year
	single point difference < + 3%


	Primary Standards usually transported to EPA Regions SRP for comparison
	

	   (if recertified via a transfer  standard)
	1/year
	Regression slopes = 1.00 + 0.03 and two intercepts are 0 + 3 ppb
	
	

	Ozone Transfer standard
	
	
	
	

	  Qualification 
	Upon receipt of transfer standard
	+4% or  +4 ppb (whichever greater)


	Transfer Standard Doc EPA 600/4-79-056 Section 6.4
	

	  Certification
	After qualification and upon receipt/adjustment/repair
	RSD of six slopes < 3.7%

Std. Dev. of  6 intercepts 1.5
	Transfer Standard Doc EPA 600/4-79-056 Section 6.6
	

	Recertification to local primary  standard
	Beginning and end of O3 season or 1/6 months whichever less
	New slope = + 0.05 of previous and
RSD of six slopes  <3.7% 

Std. Dev. of  6 intercepts 1.5
	1 recertification test that then gets added to most recent 5 tests. If does not meet acceptability certification fails
	

	Lower detectable level
	1/year
	0.003 ppm
	
	

	SYSTEMATIC CRITERIA- Ozone

	Requirement
	Frequency 
	Acceptance Criteria
	Information /Action
	

	Standard Reporting Units
	All data
	ppm (final units in AQS)
	
	

	Completeness (seasonal)
	Daily
	75% of hourly averages for the 8-hour period
	8-Hour Average
	

	Sample Residence Times
	
	< 20 seconds
	
	

	Sample Probe, Inlet, Sampling train
	
	Borosilicate glass (e.g., Pyrex®) or Teflon®
	40 CFR Part 58 App E
	

	Siting
	
	Un-obstructed probe inlet
	40 CFR Part 58 App E 
	

	EPA Standard Ozone Reference Photometer  (SRP) Recertification
	1/year
	Regression slope = 1.00 + 0.01

and intercept < 3 ppb
	This is usually at a Regional Office and is compared against the  traveling SRP 
	


CO Validation Template
	Requirement
	Frequency
	Acceptance Criteria
	Information /Action
	

	CRITICAL CRITERIA-CO

	One Point QC Check
Single analyzer
	1/ 2 weeks
	< +10% (percent difference)
	1 - 10 ppm

Relative to routine concentrations

40 CFR Part 58 App A Sec 3.2
	

	Zero/span check 


	1/ 2  weeks
	Zero drift < +  2% of full scale

Span drift < +  10 %
	
	

	OPERATIONAL CRITERIA-CO

	Shelter Temperature
	
	
	
	

	   Temperature range
	Daily

(hourly  values)
	20 to 30 C.  (Hourly ave)

or

per manufacturers specifications if designated  to a wider temperature range
	Generally the 20-30  C range will apply but  the most restrictive operable range of the instruments in the shelter may also be used as guidance
	

	   Temperature Control
	Daily (hourly values)
	< +  2 C SD over 24 hours
	
	

	   Temperature Device Check
	2/year
	+ 2o C of standard
	
	

	Precision(using 1-point QC checks)
	Calculated annually and as appropriate for design value estimates
	90% CL CV < 10%
	90% Confidence Limit of coefficient of variation.  40 CFR Part 58 App A sec 4.1.2
	

	Bias (using 1-point QC checks)
	Calculated annually and as appropriate for design value estimates
	95% CL <  + 10%
	95% Confidence Limit of absolute bias estimate 40 CFR Part 58 App A sec 4.1.3
	

	Annual Performance Evaluation
	
	
	
	

	   Single analyzer
	Every site 1/year  25 % of sites quarterly


	Percent difference of  audit levels 3-10 <  +15% 
Audit levels 1&2 + 0.03 ppm difference or  +15%
	3 consecutive audit concentration not including zero. 40 CFR Part 58 App A sec 3.2.2
	

	   Primary QA Organization 

  (PQAO)
	annually
	95% of audit percent differences fall within the one point QC check 95% probability intervals at PQAO level of aggregation
	40 CFR Part 58 App A sec 4.1.4
	

	Federal Audits (NPAP)
	1/year at selected sites 20% of sites audited
	Mean absolute difference <  15%
	40 CFR Part 58 App A sec 2.4
	

	State audits
	1/year
	State requirements
	
	

	Verification/Calibration
	Upon receipt/adjustment/repair/ installation/moving

1/6 months if manual zero/span performed biweekly

1/year if continuous zero/span performed daily
	All points within + 2 % of full scale of best-fit straight line
	Multi-point calibration

  (0 and 4 upscale points)
	

	Gaseous Standards
	
	NIST Traceable

(e.g., EPA Protocol Gas)
	Vendor must participate in EPA Protocol Gas Verification Program 40 CFR Part  58 App A sec 2.6.1
	

	Zero Air/Zero Air Check
	1/year
	Concentrations below LDL
	
	

	Gas Dilution Systems
	1/3 months
	Accuracy + 2 %
	Need to determine what this means
	

	Detection
	
	
	
	

	   Noise
	NA
	0.50 ppm
	40 CFR  Part 53.20
	

	   Lower detectable level
	1/year
	1.0 ppm
	40 CFR  Part 53.20
	

	SYSTEMATIC CRITERIA-CO

	Standard Reporting Units
	All data
	ppm (final units in AQS)
	
	

	Completeness (seasonal)
	Hourly
	75%  of hourly averages for the 8-hour period
	8-Hour average
	

	Sample Residence Times
	
	< 20 seconds
	
	

	Sample Probe, Inlet, Sampling train
	
	Borosilicate glass (e.g., Pyrex®) or Teflon® 
	40 CFR Part 58 App E 
	

	Siting
	
	Un-obstructed probe inlet
	40 CFR Part 58 App E 
	


NO2 Validation Template
	Requirement
	Frequency
	Acceptance Criteria
	Information /Action
	

	CRITICAL CRITERIA- NO2

	One Point QC Check
Single analyzer
	1/ 2 weeks
	< +10% (percent difference)
	0.01 - 0.10 ppm

Relative to routine concentrations

40 CFR Part 58 App A Sec 3.2
	

	Zero/span check 


	1/ 2  weeks
	Zero drift < +3% of full scale

Span drift < + 10 %
	
	

	OPERATIONAL CRITERIA- NO2

	Shelter Temperature
	
	
	
	

	   Temperature range
	Daily

(hourly  values)
	20 to 30 C.  (Hourly ave)

or

per manufacturers specifications if designated  to a wider temperature range
	Generally the 20-30  C range will apply but  the most restrictive operable range of the instruments in the shelter may also be used as guidance
	

	   Temperature Control
	Daily (hourly values)
	< + 2o C SD over 24 hours
	
	

	   Temperature Device Check
	2/year
	+ 2o C of standard
	
	

	Precision (using 1-point QC checks)
	Calculated annually and as appropriate for design value estimates
	90% CL CV < 10%
	90% Confidence Limit of coefficient of variation. 40 CFR Part 58 App A sec 4.1.2
	

	Bias (using 1-point QC checks)
	Calculated annually and as appropriate for design value estimates
	95% CL <  + 10%
	95% Confidence Limit of absolute bias estimate.  40 CFR Part 58 App A sec 4.1.3
	

	Annual Performance Evaluation
	
	
	
	

	   Single analyzer
	Every site 1/year  25 % of sites quarterly


	Percent difference of  audit levels 3-10 <  +15% 
Audit levels 1&2 + 1.5 ppb difference or  +15%
	3 consecutive audit concentration not including zero. 40 CFR Part 58 App A sec 3.2.2
	

	   Primary QA Organization  

   (PQAO) 
	annually
	95% of audit percent differences fall within the one point QC check 95% probability intervals  at PQAO level of aggregation
	40 CFR Part 58 App A sec 4.1.4
	

	Federal Audits (NPAP)
	1/year at selected sites 20% of sites audited
	Mean absolute difference  <15%
	40 CFR Part 58 App A sec 2.4
	

	State audits
	1/year
	State requirements
	
	

	Verification/Calibration
	Upon receipt/adjustment/repair/ installation/moving

1/6 months if manual zero/span performed biweekly

1/year if continuous zero/span performed daily
	Intrument residence time < 2 min

Dynam. parameter > 2.75 ppm-min

All points within  +2 % of full scale of best-fit straight line
	Multi-point calibration  (0 and 4 upscale points) 40 CFR Part 50 App F
	

	Converter Efficiency
	During multi-point calibrations, span and audit

1/ 2 weeks
	96%
	
	

	Gaseous Standards
	
	NIST Traceable

(e.g., EPA Protocol Gas)
	Vendor must participate in EPA Protocol Gas Verification Program 40 CFR Part  58 App A sec 2.6.1
	

	Zero Air/ Zero Air Check
	1/year
	Concentrations below LDL
	
	

	Gas Dilution Systems
	1/3 months
	Accuracy + 2 %
	Need to determine what this means
	

	Detection
	
	
	
	

	Noise
	NA
	0.005 ppm
	40 CFR  Part 53.20
	

	Lower detectable level
	1/year
	0.01 ppm
	40 CFR  Part 53.20
	

	SYSTEMATIC CRITERIA- NO2

	Standard Reporting Units
	All data
	ppm (final units in AQS)
	
	

	Completeness (seasonal)
	Quarterly
	75% 
	Annual standard (hourly data)
	

	Sample Residence Times
	
	< 20 seconds
	
	

	Sample Probe, Inlet, Sampling train
	
	Borosilicate glass (e.g., Pyrex®) or Teflon®
	40 CFR Part 58 App E
	

	Siting
	
	Un-obstructed probe inlet
	40 CFR Part 58 App E 
	


SO2  Validation Template 

	Requirement
	Frequency
	Acceptance Criteria
	Information /Action
	

	CRITICAL CRITERIA- SO2

	One Point QC Check
Single analyzer
	1/ 2 weeks
	< +10% (percent difference)
	0.01 - 0.10 ppm

Relative to routine concentrations

40 CFR Part 58 App A Sec 3.2
	

	Zero/span check 


	1/ 2  weeks
	Zero drift < + 3% of full scale

Span drift < + 10 %
	
	

	OPERATIONAL CRITERIA- SO2

	Shelter Temperature
	
	
	
	

	   Temperature range
	Daily

(hourly  values)
	20 to 30 C.  (Hourly ave)

or

per manufacturers specifications if designated  to a wider temperature range
	Generally the 20-30  C range will apply but  the most restrictive operable range of the instruments in the shelter may also be used as guidance
	

	   Temperature Control
	Daily (hourly values)
	< + 2o C SD over 24 hours
	
	

	   Temperature Device Check
	2/year
	+ 2o C  of standard
	
	

	Precision (using 1-point QC checks)
	Calculated annually and as appropriate for design value estimates
	90% CL CV < 10%
	90% Confidence Limit of coefficient of variation 40 CFR Part 58 App A sec 4.1.2
	

	Bias (using 1-point QC checks)
	Calculated annually and as appropriate for design value estimates
	95% CL <  + 10%
	95% Confidence Limit of absolute bias estimate 40 CFR Part 58 App A sec 4.1.3
	

	Annual Performance Evaluation
	
	
	
	

	   Single analyzer
	Every site 1/year  25 % of sites quarterly


	Percent difference of  audit levels 3-10 <  +15% 
Audit levels 1&2 + 1.5 ppb  difference or  +15%
	3 consecutive audit concentrations not including zero 40 CFR Part 58 App A sec 3.2.2
	

	   Primary QA Organization 

   (PQAO)
	annually
	95% of audit percent differences fall within the one point QC check 95% probability intervals  at PQAO level of aggregation
	40 CFR Part 58 App A sec 4.1.4
	

	Federal Audits (NPAP)
	1/year at selected sites 20% of sites audited
	Mean absolute difference + 15%
	40 CFR Part 58 App A sec 2.4
	

	State audits
	1/year
	State requirements
	
	

	Verification/Calibration
	Upon receipt/adjustment/repair/ installation/moving

1/6 months if manual zero/span performed biweekly

1/year if continuous zero/span performed daily
	All points within +  2 % of full scale of best-fit straight line
	Multi-point calibration

  (0 and 4 upscale points)
	

	Zero Air
	
	Concentrations below LDL
	
	

	Gaseous Standards
	
	NIST Traceable

(e.g., EPA Protocol Gas)
	Vendor must participate in EPA Protocol Gas Verification Program 40 CFR Part  58 App A sec 2.6.1
	

	Zero Air/ Zero Air Check
	1/year
	Concentrations below LDL
	
	

	Gas Dilution Systems
	1/3 months
	Accuracy + 2 %
	Need to determine what this means
	

	Detection
	
	
	
	

	   Noise
	NA
	0.005 ppm
	40 CFR  Part 53.20
	

	   Lower detectable level
	1/year
	0.01 ppm
	40 CFR  Part 53.20
	

	SYSTEMATIC CRITERIA- SO2

	Standard Reporting Units
	All data
	ppm (final units in AQS)
	
	

	Completeness (seasonal)
	Quarterly
	75% 
	Annual standard
	

	
	24 hours
	75%
	24-hour standard
	

	
	3 hours
	75%
	3-hour standard
	

	Sample Residence Times
	
	< 20 seconds
	
	

	Sample Probe, Inlet, Sampling train
	
	Borosilicate glass (e.g., Pyrex®) or Teflon® 
	40 CFR Part 58 App E
	

	Siting
	
	Un-obstructed probe inlet
	40 CFR Part 58 App E  
	


PM2.5 Filter Based Local Conditions Validation Template 

	Criteria
	Frequency
	Acceptable Range
	Information (CFR or Method 2.12)
	

	CRITICAL CRITERIA- PM2.5 Filter Based Local Conditions

	Filter Holding Times
	
	
	
	

	Sample Recovery
	all filters
	 7 days 9 hours from sample end date
	Part 50 App L Sec 10.10
	

	Post-sampling Weighing
	all filters
	 10 days from sample end date if shipped at ambient temp, or

 30 days  if shipped below ave ambient (or 4 C  or below for ave sampling temps < 4 C )  from sample end date
	Part 50 App L Sec  8..3.6
	

	Sampling Period (including multiple power failures)
	all filters
	1380-1500 minutes, or

value if < 1380 and exceedance of NAAQS 1/
midnight to midnight 
	Part 50 App L Sec  3.3

Part 50, App.L Sec 7.4.15
	

	Sampling Instrument
	
	
	
	

	Average Flow Rate
	every 24 hours of op
	average within 5% of 16.67 liters/minute
	Part 50 App L Sec 7.4
	

	Variability in Flow Rate
	every 24 hours of op
	CV   2%
	Part 50, App.L Sec 7.4.3.2
	

	Filter
	
	
	
	

	Visual Defect Check (unexposed)
	all filters
	see reference
	Part 50, App.L Sec 10.2
	

	Filter Conditioning Environment
	
	
	
	

	Equilibration
	all filters
	24 hours minimum
	Part 50, App.L Sec 8.2
	

	Temp. Range
	all filters
	24-hr mean 20-23o C
	Part 50, App.L Sec 8.2
	

	Temp.Control
	all filters
	 2o C SD* over 24 hr
	Part 50, App.L Sec 8.2
	

	Humidity Range
	all filters
	24-hr mean 30% - 40% RH or

 5% sampling RH but > 20%RH
	Part 50, App.L Sec 8.2
	

	Humidity Control
	all filters
	 5% SD* over 24 hr.
	Part 50, App.L Sec 8.2
	

	Pre/post Sampling RH
	all filters
	difference in 24-hr means    5% RH
	Part 50, App.L Sec 8.3.3
	

	Balance
	all filters
	located in filter conditioning environment
	Part 50, App.L Sec 8.3.2
	

	Verification/Calibration
	
	
	
	

	One-point Flow Rate Verification
	1/4 weeks
	 4% of transfer standard
	Part 50, App.L, Sec 9.2.5

Part 58, Appendix A Sec 3.2.3 & 3.3.2
	

	OPERATIONAL EVALUATIONS TABLE PM2.5 Filter Based Local Conditions

	Filter Checks 
	
	
	
	

	Lot Blanks
	9 filters per lot
	less than 15 µg change between weighings
	Method 2.12 Sec. 7.7
	

	Exposure Lot Blanks
	3 filters per lot
	less than 15 µg change between weighings
	Method 2.12 Sec. 7.7
	

	Filter Integrity (exposed)
	each filter
	no visual defects
	Method 2.12 Sec. 8.2
	

	Filter Holding Times
	
	
	
	

	Pre-sampling
	all filters
	< 30 days before sampling
	Part 50, App.L Sec 8.3
	

	Lab QC Checks
	
	
	
	

	Field Filter Blank
	10% or 1 per weighing session
	 30 µgchange between weighings
	Part 50, App.L Sec 8.3
	

	Lab Filter Blank
	10% or 1 per weighing session
	 15 µgchange between weighings
	Part 50, App.L Sec 8.3
	

	Balance Check
	beginning, 10th sample, end
	3 µg
	Method Sec. 7.9
	

	Duplicate Filter Weighing
	1 per weighing session
	 15 µgchange between weighings
	Method Sec 7.11
	

	Sampling Instrument
	
	
	
	

	Individual Flow Rates
	every 24 hours of op
	no flow rate excursions > 5% for > 5 min. 1/
	Part 50, App.L Sec 7.4.3.1
	

	Filter Temp Sensor
	every 24 hours of op
	no excursions of > 5 C lasting longer than 30 min 1/
	Part 50, App.L Sec 7.4
	

	Verification/Calibration
	
	
	
	

	Routine Verifications
	
	
	
	

	External Leak Check
	every 5 sampling events
	< 80 mL/min
	Part 50, App.L, Sec 7.4
	

	Internal Leak Check
	every 5 sampling events
	< 80 mL/min
	Part 50, App.L, Sec 7.4
	

	One-point Temp Verification
	1/4 weeks
	 2oC 
	Part 50, App.L, Sec 9.3
	

	Pressure Verification
	1/4 weeks
	 10 mm Hg
	Part 50, App.L, Sec 9.3
	

	Lab Temperature
	1/6 months
	 2oC
	Method Sec 3.3
	

	Lab Humidity
	1/6 months
	 2%
	Method Sec 3.3
	

	Annual Multi-point Verifications /Calibrations
	
	
	
	

	Temperature  multi-point Verification/Calibration
	1/yr
	 2oC 
	Part 50, App.L, Sec 9.3
	

	Pressure  Verification/Calibration
	on installation, then 1/yr
	 10 mm Hg
	Part 50, App.L, Sec 9.3
	

	Flow Rate  Multi-point Verification/ Calibration
	1/yr
	 2% of transfer standard
	Part 50, App.L, Sec 9.2
	

	Design Flow Rate Adjustment
	at one-point or multi-point
	 2% of design flow rate
	Part 50, App.L, Sec 9.2.6
	

	Other Monitor Calibrations
	per manufacturers’ op manual
	per manufacturers’ operating manual
	
	

	Microbalance Calibration
	1/yr
	Manufacturer’s specification
	Part 50, App.L, Sec 8.1
	

	Precision
	
	
	
	

	Collocated Samples
	every 12 days for 15% of sites
	CV < 10% of samples > 3 µg/m3
	Part 58 App A Sec 3.2.5
	

	Accuracy
	
	
	
	

	Temperature Audit
	2/yr  (method 2.12 4/yr)
	 2oC
	Method Sec. 10.2
	

	Pressure Audit
	2/yr (method 2.12 4/yr)
	10 mm Hg
	Method Sec. 10.2
	

	Microbalance Audit
	1/yr
	 0.050 mg or manufacturers specs, whichever is tighter
	Method Sec. 10.2
	

	Semi Annual Flow Rate Audit
	2/yr 
	 4% of audit standard 

 5% of design flow rate
	Part 58, App A, Sec 3.3.3
	

	Calibration & Check Standards - 
	

	Field Thermometer
	1/yr
	 0.1o C resolution,  0.5o C accuracy
	Method Sec 4.2 & 6.4
	

	Field Barometer
	1/yr
	 1 mm Hg  resolution,  5 mm Hg accuracy
	Method Sec 4.2 & 6.5
	

	Working Mass Stds. (compare to primary standards)
	1/3 mo.
	0.025 mg
	Method Sec 4.3 and 7.3
	

	Monitor Maintenance
	
	
	
	

	Impactor (WINs)

Very Sharp Cut Cyclone
	every 5 sampling events

Every 30 days
	cleaned/changed
	Method Sec 9.2
	

	Inlet/downtube Cleaning
	every 15 sampling events
	cleaned
	Method Sec 9.3
	

	Filter Chamber Cleaning
	1/4 weeks
	cleaned
	Method Sec 9.3
	

	Leak Check @
	
	see Verification/Calibration
	
	

	Circulating Fan Filter Cleaning
	1/4 weeks
	cleaned/changed
	Method Sec 9.3
	

	Manufacturer-Recommended Maintenance
	per manufacturers’ SOP
	per manufacturers’ SOP
	
	

	SYSTEMATIC CRITERIA -PM2.5  Filter Based Local Conditions

	Data Completeness
	quarterly
	> 75%
	Part 50, App. N, Sec. 4.1 (b) 4.2 (a)
	

	Reporting Units
	all filters
	µg/m3 at ambient temp/pressure (PM2.5)
	Part 50.3
	

	Rounding Convention
	
	
	
	

	Annual 3-yr average
	quarterly
	nearest 0.1 µg/m3 (> 0.05 round up)
	Part 50, App. N Sec 2.3
	

	24-hour, 3-year average
	quarterly
	nearest 1 µg/m3 (> 0.5 round up)
	Part 50, App. N Sec 2.3
	

	Detection Limit
	
	
	
	

	Lower DL
	all filters
	2 µg/m3
	Part 50, App.L Sec 3.1
	

	Upper Conc. Limit
	all filters
	 200 µg/m3
	Part 50, App.L Sec 3.2
	

	Verification/Calibration Standards Recertifications – All standards should have multi-point certifications against NIST Traceable standards
	

	Flow Rate Transfer Std.
	1/yr
	 2% of NIST-traceable Std.
	Part 50, App.L Sec 9.1 & 9.2
	

	Field Thermometer
	1/yr
	 0.1 C resolution,  0.5o C accuracy
	Method Sec 4.2.2
	

	Field Barometer
	1/yr
	 1 mm Hg resolution,  5 mm Hg accuracy
	Method Sec 4.2.2
	

	Primary Mass Stds. (compare to NIST-traceable standards)
	1/yr
	0.025 mg
	Method Sec 4.3.7
	

	Microbalance
	
	
	
	

	Readability
	at purchase
	1 µg
	Part 50, App.L Sec 8.1
	

	Repeatability
	1/yr
	1 µg 
	
	

	Calibration & Check Standards
	
	
	
	

	Flow Rate Transfer Std.
	1/yr
	 2% of NIST-traceable Std.
	Part 50, APP L, Sec 9.1 & 9.2
	

	Verification/Calibration
	
	
	
	

	Clock/timer Verification
	1/4 weeks
	1 min/mo
	Part 50, App.L, Sec 7.4
	

	Precision
	
	
	
	

	Single analyzer
	1/3 mo.
	Coefficient of variation (CV) < 10%
	
	

	Single analyzer
	1/ yr
	CV < 10%
	
	

	Primary Quality Assurance Org. 
	Annual and 3 year estimates
	90% CL of CV <  10%
	Part 58, App A, Sec 4.3.1
	

	Bias
	
	
	
	

	Performance Evaluation Program (PEP)
	5 audits for PQAOs with < 5 sites 

8 audits for PQAOs with > 5 sites
	10%
	Part 58, App A, Sec 3.2.7, 4.3.2
	


1/     value must be flagged     * SD= standard deviation    CV= coefficient of variation  @ =  Scheduled to occur immediately after impactor cleaned/changed.
NOTE:  There may be a number of continuous monitors that may be designated as an FEM or an ARM.  These monitors may have different measurement or sampling attributes that cannot be identified in this validation template. Monitoring organizations should review specific instrument operating manuals to augment this validation template as necessary. In general, 40 CFR Part 58 App A and 40 CFR part 50 App L requirements apply to Continuous PM2.5
Continuous PM2.5  Local Conditions Validation Template 

	Criteria
	Frequency
	Acceptable Range
	Information (CFR or Method 2.12)
	

	CRITICAL CRITERIA- PM2.5  Continuous, Local Conditions

	Sampling Period  
     24 hour estimate
	every  sample period
	1380-1500 minutes, or

value if < 1380 and exceedance of NAAQS 1/
midnight to midnight 
	Part 50 App L Sec  3.3

Part 50, App.L Sec 7.4.15
	

	    Hour estimate
	Every hour
	Instrument dependent
	See operators manual
	

	Sampling Instrument
	
	
	
	

	Average Flow Rate
	every 24 hours of op
	average within 5% of 16.67 liters/minute
	Part 50 App L Sec 7.4
	

	Variability in Flow Rate
	every 24 hours of op
	CV < 2%
	Part 50, App.L Sec 7.4.3.2
	

	Verification/Calibration
	
	
	
	

	One-point Flow Rate Verification
	1/4 weeks
	 4% of transfer standard
	Part 50, App.L, Sec 9.2.5

Part 58, Appendix A Sec 3.2.3 & 3.3.2
	

	Reference Membrane Verification (BAM)
	Hourly
	 4% of ABS Value
	
	

	OPERATIONAL CRITERIA- PM2.5  Continuous, Local Conditions

	Verification/Calibration
	
	
	
	

	Leak Check
	every 30 days
	Instrument dependent
	Part 50, App.L, Sec 7.4
	

	Temperature Calibration
	if multi-point failure
	 2C 
	Part 50, App.L, Sec 9.3
	

	Temp M-point Verification
	on installation, then 1/yr
	 2C 
	Part 50, App.L, Sec 9.3
	

	One-point Temp Check
	1/4 weeks
	 2C 
	Part 50, App.L, Sec 9.3
	

	Pressure Calibration
	on installation, then 1/yr
	10 mm Hg
	Part 50, App.L, Sec 9.3
	

	Pressure Verification
	1/4 weeks
	 10 mm Hg
	Part 50, App.L, Sec 9.3
	

	Other Monitor Calibrations
	per manufacturers’ op manual
	per manufacturers’ operating manual
	
	

	Flow Rate (FR) Calibration
	if multi-point verification failure
	 2% 
	Part 50, App.L, Sec 9.2
	

	FR Multi-point Verification
	1/yr
	 2% 
	Part 50, App.L, Sec 9.2
	

	Design Flow Rate Adjustment
	at one-point or multi-point
	 2% of design flow rate
	Part 50, App.L, Sec 9.2.6
	

	Precision
	
	
	
	

	Collocated Samples
	every 12 days for 15% of sites
	CV < 10% of samples > 3 µg/m3
	Part 58 App A Sec 3.2.5
	

	Accuracy
	
	
	
	

	Temperature Audit
	2/yr
	 2C
	Method 2.12 Sec. 10.2
	

	Pressure Audit
	2/yr
	10 mm Hg
	Method 2.12 Sec. 10.2
	

	Semi Annual Flow Rate Audit
	2/yr 
	 4% of audit standard 

 5% of design flow rate
	Method 2.12 Sec. 10.2
	

	Calibration & Check Standards
(working standards)
	
	
	Part 58, App A, Sec 3.3.3
	

	Field Thermometer
	1/yr
	 0.1 C resolution,  0.5 C accuracy
	Method 2.12 Sec 4.2 & 6.4
	

	Field Barometer
	1/yr
	 1 mm Hg  resolution,  5 mm Hg accuracy
	Method 2.12 Sec 4.2 & 6.5
	

	Shelter Temperature
	
	
	
	

	   Temperature range
	Daily

(hourly  values)
	20 to 30 C.  (Hourly ave)

or

per manufacturers specifications if designated  to a wider temperature range
	Generally the 20-30  C range will apply but  the most restrictive operable range of the instruments in the shelter may also be used as guidance
	

	   Temperature Control
	Daily (hourly values)
	  2o C SD over 24 hours
	
	

	   Temperature Device Check
	2/year
	 2o C 
	
	

	Monitor Maintenance
	
	
	
	

	Virtual Impactor 

Very Sharp Cut Cyclone
	Every 30 days
	cleaned/changed
	Method 2.12 Sec 9.2
	

	Inlet  Cleaning
	Every 30 days
	cleaned
	Method 2.12 Sec 9.3
	

	Filter Chamber Cleaning
	1/4 weeks
	cleaned
	Method 2.12 Sec 9.3
	

	Circulating Fan Filter Cleaning
	1/4 weeks
	cleaned/changed
	Method 2.12 Sec 9.3
	

	Manufacturer-Recommended 
Maintenance
	per manufacturers’ SOP
	per manufacturers’ SOP
	
	

	SYSTEMATIC CRITERIA- PM2.5  Continuous, Local Conditions

	Data Completeness
	quarterly
	> 75%
	Part 50, App. N, Sec. 4.1 (b) 4.2 (a)
	

	Reporting Units
	
	µg/m3 at ambient temp/pressure (PM2.5)
	Part 50.3
	

	Rounding Convention
	
	
	
	

	Annual 3-yr average
	quarterly
	nearest 0.1 µg/m3 (> 0.05 round up)
	Part 50, App. N Sec 2.3
	

	24-hour, 3-year average
	quarterly
	nearest 1 µg/m3 (> 0.5 round up)
	Part 50, App. N Sec 2.3
	

	Detection Limit
	
	
	
	

	Lower DL
	all filters
	2 µg/m3
	Part 50, App.L Sec 3.1
	

	Upper Conc. Limit
	all filters
	> 200 µg/m3
	Part 50, App.L Sec 3.2
	

	Verification/Calibration Standards Recertifications - All standards should have multi-point certifications against NIST Traceable standards
	

	Flow Rate Transfer Std.
	1/yr
	 2% of NIST-traceable Std.
	Part 50, App.L Sec 9.1 & 9.2
	

	Field Thermometer
	1/yr
	 0.1 C resolution,  0.5o C accuracy
	Method 2.12 Sec 4.2.2
	

	Field Barometer
	1/yr
	 1 mm Hg resolution,  5 mm Hg accuracy
	Method 2.12 Sec 4.2.2
	

	Calibration & Check Standards
	
	
	
	

	Flow Rate Transfer Std.
	1/yr
	 2% of NIST-traceable Std.
	Part 50, APP L, Sec 9.1 & 9.2
	

	Verification/Calibration
	
	
	
	

	Clock/timer Verification
	1/4 weeks
	1 min/mo**
	Part 50, App.L, Sec 7.4
	

	Precision
	
	
	
	

	Single analyzer
	1/3 mo.
	Coefficient of variation (CV) < 10%
	
	

	Single analyzer
	1/ yr
	CV < 10%
	
	

	Primary Quality Assurance Org. 
	Annual and 3 year estimates
	90% CL of CV <  10%
	Part 58, App A, Sec 4.3.1
	

	Bias
	
	
	
	

	Performance Evaluation Program (PEP)
	5 audits for PQAOs with < 5 sites 

8 audits for PQAOs with > 5 sites
	10%
	Part 58, App A, Sec 3.2.7, 4.3.2
	


1/     value must be flagged due to current implementation of BAM  ( sampling 42 minute/hour) only 1008 minutes of sampling in 24 hour period
*=    not defined in CFR

SD= standard deviation

CV= coefficient of variation

@ = Scheduled to occur immediately after impactor cleaned/changed.
** = need to ensure data system stamps appropriate time period with reported sample value
NOTE: The following validation template was constructed for use of PM10 at local conditions where PM10  is used in the calculation of  the PM10-2.5 measurement or for objectives other than comparison to the PM10 NAAQS.  Although the PM 10-2.5 method is found in 40 CFR Part 50 Appendix O,  Appendix O references Appendix L (the PM2.5 Method) for the QC requirements listed below.  Monitoring organizations using PM10 data for a NAAQS comparison purposes should refer to the PM10 validation template for STP (standard temperature and pressure correction). 
PM10 Filter Based Local Conditions Validation Template
	Criteria
	Frequency
	Acceptable Range
	Information (CFR or Method 2.12)
	

	CRITICAL CRITERIA- PM10 Filter Based Local Conditions

	Filter Holding Times
	
	
	
	

	Sample Recovery
	all filters
	< 7 days 9 hours from sample end date
	Part 50 App L Sec 10.10
	

	Post-sampling Weighing
	all filters
	<  10 days from sample end date if shipped at ambient temp, or

<  30 days  if shipped below ave ambient (or 4 C  or below for ave sampling temps < 4 C )  from sample end date
	Part 50 App L Sec  8..3.6
	

	Sampling Period (including multiple power failures)
	all filters
	1380-1500 minutes, or

value if < 1380 and exceedance of NAAQS 1/
midnight to midnight 
	Part 50 App L Sec  3.3

Part 50, App.L Sec 7.4.15
	

	Sampling Instrument
	
	
	
	

	Average Flow Rate
	every 24 hours of op
	average within 5% of 16.67 liters/minute
	Part 50 App L Sec 7.4
	

	Variability in Flow Rate
	every 24 hours of op
	CV <  2%
	Part 50, App.L Sec 7.4.3.2
	

	Filter
	
	
	
	

	Visual Defect Check (unexposed)
	all filters
	see reference
	Part 50, App.L Sec 10.2
	

	Filter Conditioning Environment
	
	
	
	

	Equilibration
	all filters
	24 hours minimum
	Part 50, App.L Sec 8.2
	

	Temp. Range
	all filters
	24-hr mean 20-23o C
	Part 50, App.L Sec 8.2
	

	Temp.Control
	all filters
	 2o C SD* over 24 hr
	Part 50, App.L Sec 8.2
	

	Humidity Range
	all filters
	24-hr mean 30% - 40% RH or

<  5% sampling RH but > 20%RH
	Part 50, App.L Sec 8.2
	

	Humidity Control
	all filters
	 5% SD* over 24 hr.
	Part 50, App.L Sec 8.2
	

	Pre/post Sampling RH
	all filters
	difference in 24-hr means <   5% RH
	Part 50, App.L Sec 8.3.3
	

	Balance
	all filters
	located in filter conditioning environment
	Part 50, App.L Sec 8.3.2
	

	Verification/Calibration
	
	
	
	

	One-point Flow Rate Verification
	1/4 weeks
	 4% of transfer standard
	Part 50, App.L, Sec 9.2.5

Part 58, Appendix A Sec 3.2.3 & 3.3.2
	

	OPERATIONAL EVALUATIONS TABLE PM10 Filter Based Local Conditions

	Filter Checks 
	
	
	
	

	Lot Blanks
	9 filters per lot
	less than 15 µgchange between weighings
	Method 2.12 Sec. 7.7
	

	Exposure Lot Blanks
	3 filters per lot
	less than 15 µgchange between weighings
	Method 2.12 Sec. 7.7
	

	Filter Integrity (exposed)
	each filter
	no visual defects
	Method 2.12 Sec. 8.2
	

	Filter Holding Times
	
	
	
	

	Pre-sampling
	all filters
	< 30 days before sampling
	Part 50, App.L Sec 8.3
	

	Lab QC Checks
	
	
	
	

	Field Filter Blank
	10% or 1 per weighing session
	 30 µgchange between weighings
	Part 50, App.L Sec 8.3
	

	Lab Filter Blank
	10% or 1 per weighing session
	 15 µgchange between weighings
	Part 50, App.L Sec 8.3
	

	Balance Check
	beginning, 10th sample, end
	3 µg
	Method Sec. 7.9
	

	Duplicate Filter Weighing
	1 per weighing session
	 15 µgchange between weighings
	Method Sec 7.11
	

	Sampling Instrument
	
	
	
	

	Individual Flow Rates
	every 24 hours of op
	no flow rate excursions > 5% for > 5 min. 1/
	Part 50, App.L Sec 7.4.3.1
	

	Filter Temp Sensor
	every 24 hours of op
	no excursions of > 5o C lasting longer than 30 min 1/
	Part 50, App.L Sec 7.4
	

	Verification/Calibration
	
	
	
	

	Routine Verifications
	
	
	
	

	External Leak Check
	every 5 sampling events
	< 80 mL/min
	Part 50, App.L, Sec 7.4
	

	Internal Leak Check
	every 5 sampling events
	< 80 mL/min
	Part 50, App.L, Sec 7.4
	

	One-point Temp Verification
	1/4 weeks
	 2C 
	Part 50, App.L, Sec 9.3
	

	Pressure Verification
	1/4 weeks
	 10 mm Hg
	Part 50, App.L, Sec 9.3
	

	Lab Temperature
	1/6 months
	 2o C
	Method Sec 3.3
	

	Lab Humidity
	1/6 months
	 2%
	Method Sec 3.3
	

	Annual Multi-point Verifications /Calibrations
	
	
	
	

	Temperature  multi-point Verification/Calibration
	1/yr
	2o C 
	Part 50, App.L, Sec 9.3
	

	Pressure  Verification/Calibration
	on installation, then 1/yr
	 10 mm Hg
	Part 50, App.L, Sec 9.3
	

	Flow Rate  Multi-point Verification/ Calibration
	1/yr
	 2% of transfer standard
	Part 50, App.L, Sec 9.2
	

	Design Flow Rate Adjustment
	at one-point or multi-point
	 2% of design flow rate
	Part 50, App.L, Sec 9.2.6
	

	Other Monitor Calibrations
	per manufacturers’ op manual
	per manufacturers’ operating manual
	
	

	Mirobalance Calibration
	1/yr
	Manufacturer’s specification
	Part 50, App.L, Sec 8.1
	

	Precision
	
	
	
	

	Collocated Samples
	every 12 days for 15% of sites
	CV < 10% of samples > 3 µg/m3
	Part 58 App A Sec 3.2.5
	

	Accuracy
	
	
	
	

	Temperature Audit
	2/yr
	 2o C
	Method Sec. 10.2
	

	Pressure Audit
	2/yr
	10 mm Hg
	Method Sec. 10.2
	

	Balance Audit
	1/yr
	 0.050 mg or manufacturers specs, whichever is tighter
	Method Sec. 10.2
	

	Semi Annual Flow Rate Audit
	2/yr 
	 4% of audit standard 

 5% of design flow rate
	Part 58, App A, Sec 3.3.3
	

	Calibration & Check Standards
(working standards)
	
	
	
	

	Field Thermometer
	1/yr
	 0.1o C resolution,  0.5o C accuracy
	Method Sec 4.2 & 6.4
	

	Field Barometer
	1/yr
	 1 mm Hg  resolution, 5 mm Hg accuracy
	Method Sec 4.2 & 6.5
	

	Working Mass Stds. (compare to primary standards)
	1/3 mo.
	0.025 mg
	Method Sec 4.3 and 7.3
	

	Monitor Maintenance
	
	
	
	

	Inlet/downtube Cleaning
	every 15 sampling events
	cleaned
	Method Sec 9.3
	

	Filter Chamber Cleaning
	1/4 weeks
	cleaned
	Method Sec 9.3
	

	Leak Check @
	
	see Verification/Calibration
	
	

	Circulating Fan Filter Cleaning
	1/4 weeks
	cleaned/changed
	Method Sec 9.3
	

	Manufacturer-Recommended Maintenance
	per manufacturers’ SOP
	per manufacturers’ SOP
	
	

	SYSTEMATIC CRITERIA -PM10  Filter Based Local Conditions

	Data Completeness
	quarterly
	> 75%
	Part 50, App. N, Sec. 2.1
	

	Reporting Units
	all filters
	µg/m3 at ambient temp/pressure (PM2.5)
	Part 50.3
	

	Rounding Convention
	
	
	
	

	Annual 3-yr average
	quarterly
	nearest 0.1 µg/m3 (> 0.05 round up)
	Part 50, App. N Sec 2.3
	

	24-hour, 3-year average
	quarterly
	nearest 1 µg/m3 (> 0.5 round up)
	Part 50, App. N Sec 2.3
	

	Detection Limit
	
	
	
	

	Lower DL
	all filters
	 2 µg/m3
	Part 50, App.L Sec 3.1
	

	Upper Conc. Limit
	all filters
	 200 µg/m3
	Part 50, App.L Sec 3.2
	

	Verification/Calibration Standards Recertifications- All standards should have multi-point certifications against NIST Traceable standards
	

	Flow Rate Transfer Std.
	1/yr
	 2% of NIST-traceable Std.
	Part 50, App.L Sec 9.1 & 9.2
	

	Field Thermometer
	1/yr
	 0.1o C resolution,  0.5 C accuracy
	Method Sec 4.2.2
	

	Field Barometer
	1/yr
	 1 mm Hg resolution,  5 mm Hg accuracy
	Method Sec 4.2.2
	

	Primary Mass Stds. (compare to NIST-traceable standards)
	1/yr
	0.025 mg
	Method Sec 4.3.7
	

	Microbalance
	
	
	
	

	Readability
	at purchase
	1 µg
	Part 50, App.L Sec 8.1
	

	Repeatability
	1/yr
	   1 µg
	
	

	Calibration & Check Standards
	
	
	
	

	Flow Rate Transfer Std.
	1/yr
	 2% of NIST-traceable Std.
	Part 50, APP L, Sec 9.1 & 9.2
	

	Verification/Calibration
	
	
	
	

	Clock/timer Verification
	1/4 weeks
	1 min/mo
	Part 50, App.L, Sec 7.4
	

	Precision
	
	
	
	

	Single analyzer
	1/3 mo.
	Coefficient of variation (CV) < 10%
	
	

	Single analyzer
	1/ yr
	CV < 10%
	
	

	Primary Quality Assurance Org. 
	Annual and 3 year estimates
	90% CL of CV <  10%
	Part 58, App A, Sec 4.3.1
	

	Bias
	
	
	
	

	Performance Evaluation Program (PEP)
	5 audits for PQAOs with < 5 sites 

8 audits for PQAOs with > 5 sites
	10%
	Part 58, App A, Sec 3.2.7, 4.3.2
	


1/     value must be flagged

SD= standard deviation

CV= coefficient of variation

@ =  Scheduled to occur immediately after impactor cleaned/changed.
PM10 Filter Based Dichot STP Conditions Validation Template 

	Criteria
	Frequency
	Acceptable Range
	Information (CFR or Method 2.10)
	

	CRITICAL CRITERIA- PM10 Filter Based Dichot 

	Filter Holding Times
	
	
	
	

	Sample Recovery
	all filters
	ASAP
	Part 50 App J sec 9.15
	

	Sampling Period 
	all filters
	1440 minutes +  60 minutes 

midnight to midnight
	Part 50 App  J sec 7.1.5 
	

	Sampling Instrument
	
	
	
	

	Average Flow Rate
	every 24 hours of op
	average 16.67 liters/minute
	Method 2.10 sec 2.1
	

	Filter
	
	
	
	

	Visual Defect Check (unexposed)
	all filters
	see reference
	Method 2.10 sec 4.2
	

	Collection efficiency
	all filters
	99 %


	Part 50, App J sec 7.2.2


	

	Integrity
	all filters
	 5 µg/m3
	Part 50, App J sec 7.2.3
	

	Alkalinity
	all filters
	< 25.0 microequivalents/gram
	Part 50, App J sec 7.2.4
	

	Filter Conditioning Environment
	
	
	
	

	Equilibration
	all filters
	24 hours minimum
	Part 50, App.J sec 9.3
	

	Temp. Range
	all filters
	15-30o C
	Part 50, App.J sec 7.4.1
	

	Temp.Control
	all filters
	 3 C SD* over 24 hr
	Part 50, App.J sec 7.4.2
	

	Humidity Range
	all filters
	20% - 45% RH 
	Part 50, App.J sec 7.4.3
	

	Humidity Control
	all filters
	 5% SD* over 24 hr
	Part 50, App.J sec 7.4.4
	

	Pre/post Sampling RH
	all filters
	difference in 24-hr means <  5% RH
	Part 50, App.L sec 8.3.3
	

	Balance
	all filters
	located in filter conditioning environment
	Part 50, App.L sec 8.3.2
	

	Verification/Calibration
	
	
	
	

	One-point Flow Rate Verification
	1/4 weeks
	 7% of transfer standard and 10% from design
	Method 2.10 sec Table 3-1
	

	OPERATIONAL EVALUATIONS TABLE PM10 Filter Based Dichot 

	Lab QC Checks
	
	
	
	

	Balance Check
	beginning, 10th sample, end
	4 µgof true zero
 2 µgof 10 mg check weight
	Method 2 .10 sec 4.5
	

	“Standard” filter QC check
	10% 
	 20 µgchange  from original value
	Method  2.10 sec 4.5

From standard non-routine filter
	

	“Routine” duplicate weighing
	5-7 per weighing session 
	 20 µgchange  from original value
	Method  2.10 sec 4.5

From routine filter set
	

	Verification/Calibration
	
	
	
	

	System Leak Check
	During precalibration check 
	Vacuum of 10 to 15 in. with decline to 0  >60 seconds
	Method 2.10 sec 2.2.1
	

	FR Multi-point Verification/Calibration
	1/yr
	 2% 
	Part 50, App.L, sec 9.2
	

	Field Temp M-point Verification
	on installation, then 1/yr
	 2C 
	
	

	Lab Temperature
	1/6 months
	 2C
	recommendation
	

	Lab Humidity
	1/6 months
	 2%
	recommendation
	

	Microbalance Calibration
	1/yr
	Manufacturer’s specification
	
	

	Precision
	
	
	
	

	Collocated Samples
	every 12 days for 15% of sites
	CV < 10% of samples > 3 µg/m3
	Part 58 App A sec 3.2.5
	

	Audits
	
	
	
	

	Filter Weighing
	1/yr
	 20 µgchange  from original value
	Method  2.10  Table 7-1
	

	Balance Audit
	1/yr
	Observe weighing technique and check balance with ASTM  Class 1  standard
	Method  2.10  Table 7-1 section 7.2.2
	

	Semi Annual Flow Rate Audit
	2/yr 
	 4% of audit standard 

 5% of design flow rate
	Part 58, App A, sec 3.3.3
	

	Monitor Maintenance
	
	
	
	

	Impactor
	1/3 mo
	cleaned/changed
	Method  2.10  sec 6.1.2
	

	Inlet/downtube Cleaning
	1/3 mo
	cleaned
	Method  2.10  sec 6.1.2
	

	Vacuum pump
	1/yr
	Replace diaphragm and flapper valves
	Method  2.10  sec 6.1.3
	

	Manufacturer-Recommended Maintenance
	per manufacturers’ SOP
	per manufacturers’ SOP
	
	

	SYSTEMATIC CRITERIA - PM10 Filter Based Dichot  

	Data Completeness
	quarterly
	> 75%
	Part 50 App. K, sec. 2.3
	

	Reporting Units
	all filters
	µg/m3 at standard temperature and pressure (STP)
	Part 50 App K
	

	Rounding Convention
	
	
	
	

	24-hour, 3-year average
	quarterly
	nearest 10 µg/m3 (> 5 round up)
	Part 50 App K sec 1
	

	Verification/Calibration Standards  and Recertifications - All standards should have multi-point certifications against NIST Traceable standards
	

	Flow Rate Transfer Std.
	1/yr
	 2% of NIST-traceable Std.
	Part 50, App.J sec 7.3
	

	Field Thermometer
	1/yr
	 0.1 C resolution,  0.5 C accuracy
	
	

	Field Barometer
	1/yr
	 1 mm Hg resolution,  5 mm Hg accuracy
	
	

	Primary Mass Stds. (compare to NIST-traceable standards)
	1/yr
	NIST traceable

(e.g., ANSI/ASTM Class 1, 1.1 or 2)
	Method 2.10 sec 9 
	

	Microbalance
	
	
	
	

	Readability
	at purchase
	1 µg
	Method 2.10 sec 4.4 
	

	Repeatability
	1/yr
	1 µg
	Method 2.10 sec 4.4
	

	Calibration & Check Standards
	
	
	
	

	Flow Rate Transfer Std.
	1/yr
	 2% of NIST-traceable Std.
	Method 2.10 sec 9
	

	Verification/Calibration
	
	
	
	

	Clock/timer Verification
	4/year
	5 min/mo
	recommendation
	

	Precision
	
	
	
	

	Single analyzer
	1/3 mo.
	Coefficient of variation (CV) < 10%
	recommendation
	

	Single analyzer
	1/ yr
	CV < 10%
	recommendation
	

	Primary Quality Assurance Org. 
	Annual and 3 year estimates
	90% CL of CV <  10%
	Part 58, App A, sec 4.3.1
	


SD= standard deviation   CV= coefficient of variation

PM10 Filter Based High Volume (HV) STP Conditions Validation Template 
	Criteria
	Frequency
	Acceptable Range
	Information (CFR or Method 2.11)
	

	CRITICAL CRITERIA- PM10 Filter Based  Hi-Vol

	Filter Holding Times
	
	
	
	

	Sample Recovery
	all filters
	ASAP
	Part 50 App J sec 9.15
	

	Sampling Period 
	all filters
	1440 minutes +  60 minutes 

midnight to midnight
	Part 50 App  J sec 7.1.5 
	

	Sampling Instrument
	
	
	
	

	Average Flow Rate
	every 24 hours of op
	~1.13 m3/min (varies with instrument)  
	Method 2.11 
	

	Filter
	
	
	
	

	Visual Defect Check (unexposed)
	all filters
	see reference
	Method 2.10 sec 4.2
	

	Collection efficiency 
	all filters
	99 %
	Part 50, App J sec 7.2.2
	

	Integrity 
	all filters
	 5 µg/m3
	Part 50, App J sec 7.2.3
	

	Alkalinity 
	all filters
	< 25.0 microequivalents/gram
	Part 50, App J sec 7.2.4
	

	Filter Conditioning Environment
	
	
	
	

	Equilibration
	all filters
	24 hours minimum
	Part 50, App.J sec 9.3
	

	Temp. Range
	all filters
	15-30 C
	Part 50, App.J sec 7.4.1
	

	Temp.Control
	all filters
	 3 C SD* over 24 hr
	Part 50, App.J sec 7.4.2
	

	Humidity Range
	all filters
	20% - 45% RH 
	Part 50, App.J sec 7.4.3
	

	Humidity Control
	all filters
	 5% SD* over 24 hr
	Part 50, App.J sec 7.4.4
	

	Pre/post Sampling RH
	all filters
	difference in 24-hr means   5% RH
	recommendation
	

	Balance
	all filters
	located in filter conditioning environment
	recommendation
	

	Verification/Calibration
	
	
	
	

	One-point Flow Rate Verification
	1/3 mo 
	 7% of transfer standard and 10% from design
	Method 2.10 sec Table 3-1
	

	OPERATIONAL EVALUATIONS TABLE PM10 Filter Based Hi-Vol 

	Lab QC Checks
	
	
	
	

	Balance Check
	beginning, 10th sample, end
	+  0.5 mg  of true zero and +  0.5 mg  1-5 g check weight


	Method 2 .11 sec 4.5
	

	“Routine” duplicate weighing
	5-7 per weighing session 
	 2.8 mg change  from original value 
	Method  2.11 sec 4.5.3

From routine filter set
	

	Verification/Calibration
	
	
	
	

	System Leak Check
	During precalibration check 
	Auditory inspection with faceplate blocked   
	Method 2.11 sec 2.3.2
	

	FR Multi-point Verification/Calibration
	1/yr
	3 of 4 cal points within + 10% of design
	Method 2.11 sec 2.3.2
	

	Field Temp M-point Verification
	on installation, then 1/yr
	 2C 
	
	

	Lab Temperature
	1/6 months
	 2C
	recommendation
	

	Lab Humidity
	1/6 months
	 2%
	recommendation
	

	Microbalance Calibration
	1/yr
	Manufacturer’s specification
	
	

	Precision
	
	
	
	

	Collocated Samples
	every 12 days for 15% of sites
	CV < 10% of samples > 15 µg/m3 
	Part 58 App A sec 3.2.5
	

	Audits
	
	
	
	

	Filter Weighing
	1/yr
	 5 mg change  from original value
	Method  2.11  Table 7-1
	

	Balance Audit
	1/yr
	Observe weighing technique and check balance with ASTM  Class 1  standard
	Method  2.10  Table 7-1 
	

	Semi Annual Flow Rate Audit
	2/yr 
	 10% of audit standard  and design value
	Part 58, App A, sec 3.3.3 
	

	Monitor Maintenance
	
	
	
	

	Inlet/downtube Cleaning
	1/3 mo
	cleaned
	Method  2.11  sec 6
	

	Motor/housing gaskets
	1/3 mo
	Inspected replaced
	Method  2.11  sec 6
	

	Blower motor brushes
	600-1000 hours
	Replace 
	Method  2.11  sec 6
	

	Manufacturer-Recommended Maintenance
	per manufacturers’ SOP
	per manufacturers’ SOP
	
	

	SYSTEMATIC CRITERIA - PM10 Filter Based Hi-Vol  

	Data Completeness
	quarterly
	> 75%
	Part 50 App. K, sec. 2.3
	

	Reporting Units
	all filters
	µg/m3 at standard temperature and pressure (STP)
	Part 50 App K
	

	Rounding Convention
	
	
	
	

	24-hour, 3-year average
	quarterly
	nearest 10 µg/m3 (> 5 round up)
	Part 50 App K sec 1
	

	Verification/Calibration Standards  and Recertifications - All standards should have multi-point certifications against NIST Traceable standards
	

	Flow Rate Transfer Std.
	1/yr
	 2% of NIST-traceable Std.
	Part 50, App.J sec 7.3
	

	Field Thermometer
	1/yr
	 0.1 C resolution,  0.5 C accuracy
	
	

	Field Barometer
	1/yr
	 1 mm Hg resolution,  5 mm Hg accuracy
	
	

	Primary Mass Stds. (compare to NIST-traceable standards)
	1/yr
	NIST traceable

(e.g., ANSI/ASTM Class 1, 1.1 or 2)
	Method 2.11 sec 9 
	

	Microbalance
	
	
	
	

	Readability
	at purchase
	0.1 mg
	Method 2.11 sec 4.4
	

	Repeatability
	1/yr
	0.5 mg (HV)
	Method 2.11 sec 4.4
	

	Calibration & Check Standards
	
	
	
	

	Flow Rate Transfer Std.
	1/yr
	 2% of NIST-traceable Std.
	Method 2.10 sec 9
	

	Verification/Calibration
	
	
	
	

	Clock/timer Verification
	4/year
	5 min/mo
	recommendation
	

	Precision
	
	
	
	

	Single analyzer
	1/3 mo.
	Coefficient of variation (CV) < 10%
	recommendation
	

	Single analyzer
	1/ yr
	CV < 10%
	recommendation
	

	Primary Quality Assurance Org. 
	Annual and 3 year estimates
	90% CL of CV <  10%
	Part 58, App A, sec 4.3.1
	


SD= standard deviation

CV= coefficient of variation

Continuous PM10  STP Conditions Validation Template 

NOTE:  There are a number of continuous PM10 monitors that are designated as FEM.  These monitors may have different measurement or sampling attributes that cannot be identified in this validation template. Monitoring organizations should review specific instrument operating manuals to augment this validation template as necessary. In general, 40 CFR Part 58 App A and 40 CFR part 50 App J requirements apply to Continuous PM10. Since a guidance document was never developed for continuous PM10, many of the requirements reflect a combination of manual and continuous PM2.5 requirements and are therefore considered recommendations. 
	Criteria
	Frequency
	Acceptable Range
	Information (CFR or Method 2.11)
	

	CRITICAL CRITERIA- PM10 Continuous

	Sampling Period 
	all filters
	1440 minutes +  60 minutes 

midnight to midnight
	Part 50 App  J sec 7.1.5 
	

	Sampling Instrument
	
	
	
	

	Average Flow Rate
	every 24 hours of op
	Average within + 5% of design
	recommendation
	

	Verification/Calibration
	
	
	
	

	One-point Flow Rate Verification
	1/mo 
	 7% of transfer standard and 10% from design
	Part 58, App A, sec 3.2.3
	

	OPERATIONAL EVALUATIONS TABLE PM10 Continuous 

	Verification/Calibration
	
	
	
	

	System Leak Check
	During precalibration check 
	Auditory inspection with faceplate blocked   
	Method 2.11 sec 2.3.2
	

	FR Multi-point Verification/Calibration
	1/yr
	3 of 4 cal points within + 10% of design
	Method 2.11 sec 2.3.2
	

	Audits
	
	
	
	

	Semi Annual Flow Rate Audit
	1/6 mo 
	 10% of audit standard  and design value
	Part 58, App A, sec 3.2.4 
	

	Monitor Maintenance
	
	
	
	

	Inlet/downtube Cleaning
	1/3 mo
	cleaned
	Method  2.11  sec 6
	

	Motor/housing gaskets
	1/3 mo
	Inspected replaced
	Method  2.11  sec 6
	

	Blower motor brushes
	600-1000 hours
	Replace 
	Method  2.11  sec 6
	

	Manufacturer-Recommended Maintenance
	per manufacturers’ SOP
	per manufacturers’ SOP
	
	

	SYSTEMATIC CRITERIA - PM10 Continuous 

	Data Completeness
	quarterly
	> 75%
	Part 50 App. K, sec. 2.3
	

	Reporting Units
	all filters
	µg/m3 at standard temperature and pressure (STP)
	Part 50 App K
	

	Rounding Convention
	
	
	
	

	24-hour, 3-year average
	quarterly
	nearest 10 µg/m3 (> 5 round up)
	Part 50 App K sec 1
	

	Verification/Calibration Standards  and Recertifications - All standards should have multi-point certifications against NIST Traceable standards
	

	Flow Rate Transfer Std.
	1/yr
	 2% of NIST-traceable Std.
	Part 50, App.J sec 7.3
	

	Field Thermometer
	1/yr
	 0.1 C resolution,  0.5 C accuracy
	recommendation
	

	Field Barometer
	1/yr
	 1 mm Hg resolution,  5 mm Hg accuracy
	recommendation
	

	Calibration & Check Standards
	
	
	
	

	Flow Rate Transfer Std.
	1/yr
	 2% of NIST-traceable Std.
	Method 2.10 sec 9
	

	Verification/Calibration
	
	
	
	

	Clock/timer Verification
	4/year
	5 min/mo
	recommendation
	


Pb High Volume (TSP) Validation Template

Note: in 2008, the NAAQS was lowered for Pb and new monitoring rules were promulgated which allowed for the use of federal equivalent analytical methods and the use of PM10 sampling in certain circumstances.  The following information is guidance based on the current FRM which is sampling by TSP and analysis by atomic absorption. Information is this table is derived from the TSP sampling method in 40 CFR Part 50 App B,  and QA Handbook Method 2.2 (1977).  The analytical requirements/guidance is derived from 40 CFR Part 50, App G and QA Handbook Method  2.8 (1981). Monitoring for Pb based on the new NAAQS requirements will begin in calendar year 2010.  In 2009, new guidance related to analytical FEM (ICP-MS, XRF, etc.) will be developed and included as additional guidance for Pb.  Revised and/or additional Pb validation templates will be included in this section and posted on AMTIC.
	Criteria
	Frequency
	Acceptable Range
	Information (CFR or Method 2.2 or 2.8
	

	CRITICAL CRITERIA- Pb in TSP

	Filter Holding Times
	
	
	
	

	Sample Recovery
	all filters
	ASAP
	Part 50 App B 
	

	Sampling Period 
	all filters
	1440 minutes +  60 minutes 

midnight to midnight
	Part 50 App  B sec 8.15 
	

	Sampling Instrument
	
	
	
	

	Average Flow Rate
	every 24 hours of op
	1.1-1.70 m3/min (varies with instrument)  
	Part 50 App B  sec 8.8 
	

	Filter
	
	
	Part 50 App B sec 7.1
	

	Visual Defect Check (unexposed)
	all filters
	see reference
	Part 50 App B sec 8.2
	

	Collection Efficiency 
	all filters
	99 %
	Part 50 App B sec 7.1.4
	

	Pressure Drop Range 
	all filters
	42-54 mm Hg
	Part 50 App B sec 7.1.5
	

	pH 
	all filters
	6-10
	Part 50, App B sec 7.1.6
	

	Pb Content 
	all filters pre-sampling batch check
	<75 µg/filter
	Part 50, App G sec 6.1.1

Method 2.8 sec 6.2.1
	

	Verification/Calibration
	
	
	
	

	One-point Flow Rate Verification
	1/3 mo 
	+7% from design transfer standard +10% from design
	Part 58 App A Method 2.2 sec 2.6 
	

	Calibration Reproducibility Checks
	Beginning, every 10 samples and end
	+ 5%  of value predicted by calibration curve
	Part 50, App G Sec 9.3

May be FEM dependent
	

	Reagent Blank
	Every analytical batch
	< LDL
	recomendation
	

	Daily Calibration
	Daily 
	until good agreement is obtained among replicates
	Method 2.8 sec 2.8.5
	

	OPERATIONAL EVALUATIONS TABLE Pb in TSP 

	Verification/Calibration
	
	
	
	

	System Leak Check
	During precalibration check 
	Auditory inspection with faceplate blocked   
	Method 2.2 sec 2.0
	

	FR Multi-point Verification/Calibration
	After receipt, after motor maintenance or failure of 1-point check and 
1/yr
	5 points  over  range of 1.1 to 1.7 m3/min 

within + 5% limits of linearity
	Method 2.2 sec 2.6
	

	Precision
	
	
	
	

	Collocated Samples
	15% of each method code in PQAO

Frequency - every 12 days 
	CV <  20% of samples > 0.02 µg/m3 (cutoff value)
	Part 58 App A sec 3.2.5
	

	Audits
	
	
	
	

	Semi Annual Flow Rate Audit
	2/yr 
	 10% of audit standard  and design value
	Part 58, App A, sec 3.3.4 
	

	Analysis Audits
	6 strips/quarter

3 at each concentration range
	10% (percent difference)
	Part 58, App A, sec 3.3.4
	

	Blanks
	
	
	
	

	Field Filter Blank
	1/quarter
	< LDL
	recommendation
	

	Monitor Maintenance
	
	
	
	

	Inlet cleaning
	1/3 mo
	cleaned
	recommendation
	

	Motor/housing gaskets
	~400 hours
	Inspected replaced
	Method 2.2 sec 7
	

	Blower motor brushes
	400-500
	Replace 
	Method 2.2 sec 7
	

	Manufacturer-Recommended Maintenance
	per manufacturers’ SOP
	per manufacturers’ SOP
	NA
	

	SYSTEMATIC CRITERIA - Pb Filter Based Hi-Vol  

	Data Completeness
	quarterly
	three -month mean (i.e., the 3-month data capture  rate)   > 75%
	Part 50 App. R, sec. 4
	

	Reporting Units
	all filters
	µg/m3 at local temperature and pressure.
	Part 50 App R
	

	Rounding Convention
	
	
	
	

	3-month  arithmetic mean 
	quarterly
	Report data to 3 decimal places (data after 3 are truncated.
	Part 50 App R
	

	Lower Detectable Limit
	
	
	
	

	Atomic Absorption 
	
	0.07 µg/m3
	Part 50 App G sec 2.3
	

	Verification/Calibration Standards and Recertifications - All standards should have multi-point certifications against NIST Traceable standards
	

	Flow Rate Transfer Std.
	1/yr
	Resolution 0.02 m3/min

 2% reproducibility 
	Part 50, App.B sec 7.8
	

	Field Thermometer
	1/yr
	2 C resolution
	Part 50, App.B sec 7.5
	

	Field Barometer
	1/yr
	 5 mm Hg resolution
	Part 50, App.B sec 7.6
	

	Analytical Standards
	
	
	
	

	Reagents (HNO3 and HCL)
	
	ACS reagent grade
	Part 50 App G sec.6.2
	

	Pb nitrate  Pb (NO3)2
	
	ACS reagent grade (99.0% purity
	Part 50 App G sec.6.2
	

	Verification/Calibration
	
	
	
	

	Clock/timer Verification
	4/year
	5 min/mo
	recommendation
	

	
	
	
	
	

	Precision
	
	
	
	

	Single analyzer
	1/3 mo.
	Coefficient of variation (CV) < 20%
	recommendation
	

	Single analyzer
	1/ yr
	CV < 20%
	recommendation
	

	Primary Quality Assurance Org. 
	Annual and 3 year estimates
	90% CL of CV <  20%
	Part 58, App A, sec 4.3.1
	

	Bias
	
	
	
	

	Performance Evaluation Program (PEP)
	5 audits for PQAOs with < 5 sites 

8 audits for PQAOs with > 5 sites
	95% CL Absolute bias 15%
	Part 58, App A, Sec 2.3.1
	


SD= standard deviation

CV= coefficient of variation


